## Exercise 1a
Description: A single aspirated temperature sensor reading in °C is taking measurements once a minute. For your analysis, you are using the minute readings for trending. The calibration certificate for the temperature sensor is provided. Find the expanded measurement uncertainty with a confidence level of 95% that includes repeatability and the calibration uncertainty components.
Example Application: Temperature sensor reading in °C
Food for Thought: How is repeatability acquired for this example?
Data:
Date | Time | Temperature (°C)
2016-10-10 | 1:00 | 26.1
| 1:01 | 26.1145
| 1:02 | 26.1215
| 1:03 | 26.1337
| 1:04 | 26.1172
| 1:05 | 26.146
| 1:06 | 26.151
| 1:07 | 26.1673
| 1:08 | 26.1905
| 1:09 | 26.176
| 1:10 | 26.1997

## 1a Solution
Description: A single aspirated temperature sensor reading in °C is taking measurements once a minute. For your analysis, you are using the minute readings for trending. The calibration certificate for the temperature sensor is provided. Find the expanded measurement uncertainty with a confidence level of 95% that includes repeatability and the calibration uncertainty components.
Example Application: Temperature sensor reading in °C
Food for Thought: How is repeatability acquired for this example?
Answer: You can do an experiment yourself to look at standard deviation in repeatable conditions (Type A analysis) or if the calibration certificate/manufacturer provides estimates for this, you can use these estimates (Type B analysis).
Data: |  |  |  |  | From the Sensor Calibration Certificate:
Date | Time | Temperature (°C)
2016-10-10 | 1:00 | 26.1 | ± | 0.042619 | Expanded Uncertainty | 0.04
| 1:01 | 26.1145 | ± | 0.042619 | Coverage Factor | 2
| 1:02 | 26.1215 | ± | 0.042619 | Calibration Uncertainty | 0.02
| 1:03 | 26.1337 | ± | 0.042619 | Expanded Sensor Repeatability | 0.016
| 1:04 | 26.1172 | ± | 0.042619 | Coverage Factor | 2
| 1:05 | 26.146 | ± | 0.042619 | Sensor Repeatability | 0.008
| 1:06 | 26.151 | ± | 0.042619
| 1:07 | 26.1673 | ± | 0.042619 | Uncertainty Components | n | Uncertainty (°C)
| 1:08 | 26.1905 | ± | 0.042619 | Calibration Uncertainty | 100 | 0.02
| 1:09 | 26.176 | ± | 0.042619 | Repeatability | 100 | 0.008
| 1:10 | 26.1997 | ± | 0.042619 | Combined Uncertainty |  | 0.021541
[5 empty] | Degrees of Freedom | 129.8892
[5 empty] | Expansion Factor | 1.9785
[5 empty] | Expanded Uncertainty |  | 0.042619

## Exercise 1b
Description: To better your uncertainty analysis from scenario 1a, you have acquired the manufacturer's quote for the aspirated shield's impact on temperature measurement accuracy is 0.02°C. Find the expanded measurement uncertainty with a confidence level of 95% for the minute measurements that includes repeatability, aspirated shield and the calibration uncertainty components.
Example Application: Temperature sensor reading in °C
Food for Thought: Notice the sensor's calibraiton certificate is almost up, what are some options to make the measurement traceable?
Data:
Date | Time | Temperature (°C)
2016-10-10 | 1:00 | 26.1
| 1:01 | 26.1145
| 1:02 | 26.1215
| 1:03 | 26.1337
| 1:04 | 26.1172
| 1:05 | 26.146
| 1:06 | 26.151
| 1:07 | 26.1673
| 1:08 | 26.1905
| 1:09 | 26.176
| 1:10 | 26.1997

## 1b Solution
Description: To better your uncertainty analysis from scenario 1a, you have acquired the manufacturer's quote for the aspirated shield's impact on temperature measurement accuracy is 0.02°C. Find the expanded measurement uncertainty with a confidence level of 95% for the minute measurements that includes repeatability, aspirated shield and the calibration uncertainty components.
Example Application: Temperature sensor reading in °C
Food for Thought: Notice the sensor's calibraiton certificate is almost up, what are some options to make the measurement traceable?
Answer: Recalibrate the sensor at a SI traceable facility or validate the sensor with another temperature that will allow you to remain in your uncertainty requirements.
Data: |  |  |  |  | From the Aspirated Shield Accuracy Quote:
Date | Time | Temperature (°C) |  |  | Accuracy | 0.02
2016-10-10 | 1:00 | 26.1 | ± | 0.044545 | Uncertainty | 0.0066667
| 1:01 | 26.1145 | ± | 0.044545
| 1:02 | 26.1215 | ± | 0.044545 | From the Sensor Calibration Certificate:
| 1:03 | 26.1337 | ± | 0.044545
| 1:04 | 26.1172 | ± | 0.044545 | Expanded Uncertainty | 0.04
| 1:05 | 26.146 | ± | 0.044545 | Coverage Factor | 2
| 1:06 | 26.151 | ± | 0.044545 | Calibration Uncertainty | 0.02
| 1:07 | 26.1673 | ± | 0.044545 | Sensor Repeatability | 0.008
| 1:08 | 26.1905 | ± | 0.044545
| 1:09 | 26.176 | ± | 0.044545 | Uncertainty Components | n | Uncertainty (°C)
| 1:10 | 26.1997 | ± | 0.044545 | Calibration Uncertainty | 100 | 0.02
[5 empty] | Aspirated Shield | 100 | 0.0066667
[5 empty] | Repeatability | 100 | 0.008
[5 empty] | Combined Uncertainty |  | 0.022549
[5 empty] | Degrees of Freedom | 154.1089
[5 empty] | Expansion Factor | 1.9755
[5 empty] | Expanded Uncertainty |  | 0.044545

## Exercise 1c
Description: A single aspirated temperature sensor reading in °C is taking measurements once every 10 seconds. For your analysis, you are using minute averages for trending. The calibration certificate for the sensor is provided below. Find the expanded measurement uncertainty with a confidence level of 95% that includes repeatability and the calibration uncertainty components.
Example Application: Temperature sensor reading in °C
Food for Thought: How could you further reduce uncecrtainty of these measurements?
Data:
Date | Time | Temperature (°C)
2016-10-10 | 1:00 | 26.105
| 1:00 | 26.103
| 1:00 | 26.093
| 1:01 | 26.104
| 1:01 | 26.091
| 1:01 | 26.104
| 1:01 | 26.114
| 1:01 | 26.101
| 1:01 | 26.122
| 1:02 | 26.122
| 1:02 | 26.116
| 1:02 | 26.112
| 1:02 | 26.127
| 1:02 | 26.107
| 1:02 | 26.13
| 1:03 | 26.114
| 1:03 | 26.141
| 1:03 | 26.11
| 1:03 | 26.13
| 1:03 | 26.132
| 1:03 | 26.135
| 1:04 | 26.145
| 1:04 | 26.127
| 1:04 | 26.133
| 1:04 | 26.104
| 1:04 | 26.123
| 1:04 | 26.111
| 1:05 | 26.128
| 1:05 | 26.117
| 1:05 | 26.12
| 1:05 | 26.149
| 1:05 | 26.158
| 1:05 | 26.142
| 1:06 | 26.13
| 1:06 | 26.147
| 1:06 | 26.15
| 1:06 | 26.152
| 1:06 | 26.153
| 1:06 | 26.143
| 1:07 | 26.159
| 1:07 | 26.152
| 1:07 | 26.147
| 1:07 | 26.169
| 1:07 | 26.169
| 1:07 | 26.162
| 1:08 | 26.161
| 1:08 | 26.162
| 1:08 | 26.181
| 1:08 | 26.193
| 1:08 | 26.188
| 1:08 | 26.172
| 1:09 | 26.205
| 1:09 | 26.185
| 1:09 | 26.2
| 1:09 | 26.174
| 1:09 | 26.19
| 1:09 | 26.175
| 1:10 | 26.175
| 1:10 | 26.174
| 1:10 | 26.168
| 1:10 | 26.195
| 1:10 | 26.202
| 1:10 | 26.197
| 1:11 | 26.205
| 1:11 | 26.207
| 1:11 | 26.192

## 1c Solution
Description: A single aspirated temperature sensor reading in °C is taking measurements once every 10 seconds. For your analysis, you are using minute averages for trending. The calibration certificate for the sensor is provided below. Find the expanded measurement uncertainty with a confidence level of 95% that includes repeatability and the calibration uncertainty components.
Example Application: Temperature sensor reading in °C
Food for Thought: How could you further reduce uncecrtainty of these measurements?
Answer: You could increase your sample frequency to once per second to have additional readings and better estimate your average. You could also find a calibration facility that provides a better uncertainty.
Data: |  |  |  | Analysis from the Data: | [5 empty] | From the Sensor Calibration Certificate:
Date | Time | Temperature (°C) |  |  |  | Time | result |  | 95% CL
2016-10-10 | 1:00 | 26.105 | 1 | average | 26.1 | 1:00 | 26.1 | ± | 0.039997 | Expanded Uncertainty | 0.04
| 1:00 | 26.103 | 2 | standard deviation | 0.006261 |  |  |  |  | Coverage Factor | 2
| 1:00 | 26.093 | 3 | count | 6 |  |  |  |  | Calibration Uncertainty | 0.02
| 1:01 | 26.104 | 4 | standard error | 0.002556
| 1:01 | 26.091 | 5
| 1:01 | 26.104 | 6 | [6 empty] | Uncertainty Components | n | Uncertainty (°C)
| 1:01 | 26.114 | 1 | average | 26.1145 | 1:01 | 26.1145 | ± | 0.040173 | Calibration Uncertainty | 100 | 0.02
| 1:01 | 26.101 | 2 | standard deviation | 0.007791 |  |  |  |  | Repeatability | 6 | 0.002556
| 1:01 | 26.122 | 3 | count | 6 |  |  |  |  | Combined Uncertainty |  | 0.020163
| 1:02 | 26.122 | 4 | standard error | 0.0031807 |  |  |  |  | Degrees of Freedom | 101.7231
| 1:02 | 26.116 | 5 | [6 empty] | Expansion Factor | 1.9837
| 1:02 | 26.112 | 6 | [6 empty] | Expanded Uncertainty |  | 0.039997
| 1:02 | 26.127 | 1 | average | 26.1215 | 1:02 | 26.1215 | ± | 0.041106
| 1:02 | 26.107 | 2 | standard deviation | 0.013278
| 1:02 | 26.13 | 3 | count | 6
| 1:03 | 26.114 | 4 | standard error | 0.0054206
| 1:03 | 26.141 | 5
| 1:03 | 26.11 | 6
| 1:03 | 26.13 | 1 | average | 26.1337 | 1:03 | 26.1337 | ± | 0.03999
| 1:03 | 26.132 | 2 | standard deviation | 0.006186
| 1:03 | 26.135 | 3 | count | 6
| 1:04 | 26.145 | 4 | standard error | 0.0025254
| 1:04 | 26.127 | 5
| 1:04 | 26.133 | 6
| 1:04 | 26.104 | 1 | average | 26.1172 | 1:04 | 26.1172 | ± | 0.040283
| 1:04 | 26.123 | 2 | standard deviation | 0.008612
| 1:04 | 26.111 | 3 | count | 6
| 1:05 | 26.128 | 4 | standard error | 0.0035158
| 1:05 | 26.117 | 5
| 1:05 | 26.12 | 6
| 1:05 | 26.149 | 1 | average | 26.146 | 1:05 | 26.146 | ± | 0.040399
| 1:05 | 26.158 | 2 | standard deviation | 0.0094021
| 1:05 | 26.142 | 3 | count | 6
| 1:06 | 26.13 | 4 | standard error | 0.0038384
| 1:06 | 26.147 | 5
| 1:06 | 26.15 | 6
| 1:06 | 26.152 | 1 | average | 26.151 | 1:06 | 26.151 | ± | 0.039922
| 1:06 | 26.153 | 2 | standard deviation | 0.0054772
| 1:06 | 26.143 | 3 | count | 6
| 1:07 | 26.159 | 4 | standard error | 0.0022361
| 1:07 | 26.152 | 5
| 1:07 | 26.147 | 6
| 1:07 | 26.169 | 1 | average | 26.1673 | 1:07 | 26.1673 | ± | 0.04015
| 1:07 | 26.169 | 2 | standard deviation | 0.007607
| 1:07 | 26.162 | 3 | count | 6
| 1:08 | 26.161 | 4 | standard error | 0.0031056
| 1:08 | 26.162 | 5
| 1:08 | 26.181 | 6
| 1:08 | 26.193 | 1 | average | 26.1905 | 1:08 | 26.1905 | ± | 0.040792
| 1:08 | 26.188 | 2 | standard deviation | 0.011709
| 1:08 | 26.172 | 3 | count | 6
| 1:09 | 26.205 | 4 | standard error | 0.0047802
| 1:09 | 26.185 | 5
| 1:09 | 26.2 | 6
| 1:09 | 26.174 | 1 | average | 26.176 | 1:09 | 26.176 | ± | 0.040118
| 1:09 | 26.19 | 2 | standard deviation | 0.0073485
| 1:09 | 26.175 | 3 | count | 6
| 1:10 | 26.175 | 4 | standard error | 0.003
| 1:10 | 26.174 | 5
| 1:10 | 26.168 | 6
| 1:10 | 26.195 | 1 | average | 26.1997 | 1:10 | 26.1997 | ± | 0.039963
| 1:10 | 26.202 | 2 | standard deviation | 0.0059217
| 1:10 | 26.197 | 3 | count | 6
| 1:11 | 26.205 | 4 | standard error | 0.0024175
| 1:11 | 26.207 | 5
| 1:11 | 26.192 | 6

## Exercise 1d
Description: A single aspirated temperature sensor reading in ohms is taking measurements once a second. For your analysis, you are using minute averages for trending. The calibration certificates for the sensor and data acquisition system (DAS) are provided below. Find the expanded measurement uncertainty with a confidence level of 95% for 1 minute of data that includes repeatability, DAS calibration and sensor calibration uncertainty components.
Example Application: Temperature sensor reading in ohm
Food for Thought: (1) How is the DAS uncertainty quoted in ohms transitioned to degrees C? (2) What other components could be affecting uncertainty in measurement?
Data:
Date | Time | Temperature (ohm)
2016-10-10 | 1:00 | 110.0855
| 1:00 | 110.0968
| 1:00 | 110.1233
| 1:00 | 110.1062
| 1:00 | 110.0991
| 1:00 | 110.1025
| 1:00 | 110.1111
| 1:00 | 110.1049
| 1:00 | 110.0955
| 1:00 | 110.0882
| 1:00 | 110.0923
| 1:00 | 110.1108
| 1:00 | 110.0872
| 1:00 | 110.0946
| 1:00 | 110.0959
| 1:00 | 110.1109
| 1:00 | 110.1072
| 1:00 | 110.0925
| 1:00 | 110.097
| 1:00 | 110.106
| 1:00 | 110.0906
| 1:00 | 110.0942
| 1:00 | 110.1005
| 1:00 | 110.0961
| 1:00 | 110.1062
| 1:00 | 110.0873
| 1:00 | 110.0946
| 1:00 | 110.0951
| 1:00 | 110.0854
| 1:00 | 110.093
| 1:01 | 110.0979
| 1:01 | 110.0989
| 1:01 | 110.0896
| 1:01 | 110.1069
| 1:01 | 110.1154
| 1:01 | 110.1022
| 1:01 | 110.1062
| 1:01 | 110.0895
| 1:01 | 110.1027
| 1:01 | 110.1003
| 1:01 | 110.0927
| 1:01 | 110.1044
| 1:01 | 110.0941
| 1:01 | 110.1171
| 1:01 | 110.0821
| 1:01 | 110.0948
| 1:01 | 110.0982
| 1:01 | 110.1028
| 1:01 | 110.1067
| 1:01 | 110.0956
| 1:01 | 110.0999
| 1:01 | 110.1156
| 1:01 | 110.1
| 1:01 | 110.0871
| 1:01 | 110.1145
| 1:01 | 110.1093
| 1:01 | 110.1148
| 1:01 | 110.0978
| 1:01 | 110.1146
| 1:01 | 110.1007

## 1d Solution
Description: A single aspirated temperature sensor reading in ohms is taking measurements once a second. For your analysis, you are using minute averages for trending. The calibration certificates for the sensor and data acquisition system (DAS) are provided below. Find the expanded measurement uncertainty with a confidence level of 95% for 1 minute of data that includes repeatability, DAS calibration and sensor calibration uncertainty components.
Example Application: Temperature sensor reading in resistance (ohms)
Food for Thought: (1) How is the DAS uncertainty quoted in ohms transitioned to degrees C? (2) What other components could be affecting uncertainty in measurement?
Answer: (1) The partial derivative of the calibration algorithm in respect to resistance (dT/dR) is multipled by the DAS uncertainty in ohms (dR) to quantify the uncertainty in °C (dT). (2) The aspirated shield and fan could be contributing to measurement uncertainty. Sometimes the manufacturers will have estimates for these.
Data:
Date | Time | Resistance (ohm) | Temperature (°C) |  |  | From the Sensor Calibration Certificate: |  |  | From the DAS Calibration Certificate:
2016-10-10 | 1:00 | 110.0997 | 25.807 | average (ohm) | 110.1
| 1:00 | 110.1016 | 25.8119 | average (°C) | 25.8079 | Expanded Uncertainty | 0.02 |  | Expanded Uncertainty (ohm) | 0.00039
| 1:00 | 110.1013 | 25.8112 | standard deviation | 0.002114 | Coverage Factor | 2 |  | Coverage Factor | 2
| 1:00 | 110.1008 | 25.8099 | count | 60 | Calibration Uncertainty | 0.01 |  | Calibration Uncertainty (ohm) | 0.000195 | dR
| 1:00 | 110.0992 | 25.8057 | standard error | 0.00027291 | CA2 | 0.001044
| 1:00 | 110.1004 | 25.8088 |  |  | CA1 | 2.3451
| 1:00 | 110.0991 | 25.8055 |  |  | CA0 | -245.047
| 1:00 | 110.1003 | 25.8086 | [5 empty] | Calibration Uncertainty (°C) | 0.00050213
| 1:00 | 110.1003 | 25.8086
| 1:00 | 110.1007 | 25.8096 |  |  | Uncertainty Components | n | Uncertainty (°C)
| 1:00 | 110.0998 | 25.8073 |  |  | Sensor Calibration Uncertainty | 100 | 0.01
| 1:00 | 110.0995 | 25.8065 |  |  | DAS Calibration Uncertainty | 100 | 0.00050213
| 1:00 | 110.0993 | 25.806 |  |  | Repeatability | 60 | 0.00027291
| 1:00 | 110.0995 | 25.8065 |  |  | Combined Uncertainty |  | 0.010016
| 1:00 | 110.1003 | 25.8086 |  |  | Degrees of Freedom | 99.647
| 1:00 | 110.0992 | 25.8057 |  |  | Expansion Factor | 1.9842
| 1:00 | 110.1007 | 25.8096 |  |  | Expanded Uncertainty |  | 0.019875
| 1:00 | 110.0994 | 25.8063
| 1:00 | 110.1013 | 25.8112
| 1:00 | 110.0995 | 25.8065
| 1:00 | 110.0999 | 25.8076
| 1:00 | 110.0994 | 25.8063 | From Exercise 1a with Exercise 1c measurement uncertainty
| 1:00 | 110.0994 | 25.8063
| 1:00 | 110.0999 | 25.8076 |  |  |  |  | 50 | 119.461 | 50.0041
| 1:00 | 110.1004 | 25.8088 |  |  |  |  | 0 | 100.037 | 0.0010969
| 1:00 | 110.1 | 25.8078 |  |  |  |  | -30 | 88.232 | -30.0035
| 1:00 | 110.1009 | 25.8101
| 1:00 | 110.1007 | 25.8096
| 1:00 | 110.0993 | 25.806
| 1:00 | 110.0997 | 25.807
| 1:01 | 110.0994 | 25.8063
| 1:01 | 110.1001 | 25.8081
| 1:01 | 110.1009 | 25.8101
| 1:01 | 110.1009 | 25.8101
| 1:01 | 110.1001 | 25.8081
| 1:01 | 110.0973 | 25.8009
| 1:01 | 110.1008 | 25.8099
| 1:01 | 110.0997 | 25.807
| 1:01 | 110.0985 | 25.8039
| 1:01 | 110.0993 | 25.806
| 1:01 | 110.0996 | 25.8068
| 1:01 | 110.0999 | 25.8076
| 1:01 | 110.0999 | 25.8076
| 1:01 | 110.1007 | 25.8096
| 1:01 | 110.0992 | 25.8057
| 1:01 | 110.1003 | 25.8086
| 1:01 | 110.101 | 25.8104
| 1:01 | 110.1015 | 25.8117
| 1:01 | 110.1014 | 25.8114
| 1:01 | 110.1014 | 25.8114
| 1:01 | 110.0995 | 25.8065
| 1:01 | 110.0995 | 25.8065
| 1:01 | 110.1008 | 25.8099
| 1:01 | 110.0996 | 25.8068
| 1:01 | 110.0995 | 25.8065
| 1:01 | 110.1012 | 25.8109
| 1:01 | 110.1002 | 25.8083
| 1:01 | 110.0991 | 25.8055
| 1:01 | 110.0997 | 25.807
| 1:01 | 110.0999 | 25.8076
𝑑𝑇/𝑑𝑅=2𝐶_𝐴2 (𝑅)+𝐶_𝐴1

## Exercise 1e
Description: Your temperature sensor is out of calibration, but you have acquired a temperature sensor that has just recently been calibrated with the below calibration certificate. Find the new expanded measurement uncertainty with a confidence level of 95% for the sensor out of calibration.
Example Application: Temperature sensor reading in °C
Food for Thought: (1) What additional data would you want to collect? (2) What is another component of uncertainty that would be important to consider between calibration or validation cycles?
Data:
Date | Time | Out-of-Cal Sensor (°C) | Validation Sensor (°C)
2016-10-14 | 1:00 | 25.1788 | 25.1766
| 1:00 | 25.1785 | 25.1769
| 1:00 | 25.1783 | 25.1765
| 1:00 | 25.1787 | 25.1769
| 1:00 | 25.1783 | 25.1767
| 1:00 | 25.1779 | 25.1761
| 1:00 | 25.1775 | 25.1763
| 1:00 | 25.1773 | 25.1758
| 1:00 | 25.1768 | 25.1752
| 1:00 | 25.1764 | 25.1749
| 1:00 | 25.1763 | 25.1752
| 1:00 | 25.1763 | 25.1753
| 1:00 | 25.1759 | 25.1749
| 1:00 | 25.176 | 25.1751
| 1:00 | 25.1758 | 25.1748
| 1:00 | 25.1757 | 25.1747
| 1:00 | 25.1756 | 25.1748
| 1:00 | 25.1752 | 25.1752
| 1:00 | 25.1757 | 25.1748
| 1:00 | 25.1755 | 25.1746
| 1:00 | 25.1753 | 25.1748
| 1:00 | 25.1751 | 25.1741
| 1:00 | 25.1748 | 25.175
| 1:00 | 25.1751 | 25.175
| 1:00 | 25.1754 | 25.1752
| 1:00 | 25.1756 | 25.1752
| 1:00 | 25.1754 | 25.1752
| 1:00 | 25.1756 | 25.1751
| 1:00 | 25.1752 | 25.1747
| 1:00 | 25.1749 | 25.175
| 1:01 | 25.1752 | 25.1751
| 1:01 | 25.1752 | 25.1752
| 1:01 | 25.175 | 25.1748
| 1:01 | 25.175 | 25.1746
| 1:01 | 25.1747 | 25.1747
| 1:01 | 25.175 | 25.1746
| 1:01 | 25.1749 | 25.1743
| 1:01 | 25.1745 | 25.1745
| 1:01 | 25.1744 | 25.174
| 1:01 | 25.1743 | 25.1742
| 1:01 | 25.1741 | 25.1744
| 1:01 | 25.1736 | 25.1741
| 1:01 | 25.1736 | 25.174
| 1:01 | 25.1741 | 25.1745
| 1:01 | 25.1742 | 25.174
| 1:01 | 25.1737 | 25.1743
| 1:01 | 25.1741 | 25.1748
| 1:01 | 25.1745 | 25.1745
| 1:01 | 25.1742 | 25.1751
| 1:01 | 25.1743 | 25.1755
| 1:01 | 25.1745 | 25.1755
| 1:01 | 25.1749 | 25.1754
| 1:01 | 25.1747 | 25.1752
| 1:01 | 25.1746 | 25.1759
| 1:01 | 25.175 | 25.176
| 1:01 | 25.1751 | 25.1758
| 1:01 | 25.1756 | 25.1761
| 1:01 | 25.1754 | 25.1763
| 1:01 | 25.1756 | 25.1755
| 1:01 | 25.1755 | 25.176
| 1:01 | 25.1751 | 25.1759
| 1:01 | 25.1752 | 25.1763
| 1:01 | 25.1753 | 25.1762
| 1:01 | 25.1756 | 25.1768
| 1:01 | 25.1759 | 25.1763
| 1:01 | 25.1755 | 25.1766
| 1:01 | 25.1761 | 25.1767
| 1:01 | 25.1763 | 25.1769
| 1:01 | 25.1764 | 25.1774
| 1:01 | 25.1761 | 25.1771
| 1:01 | 25.1764 | 25.1775
| 1:01 | 25.1765 | 25.177
| 1:01 | 25.177 | 25.1776
| 1:01 | 25.1765 | 25.178
| 1:01 | 25.1767 | 25.1775
| 1:01 | 25.177 | 25.1774
| 1:01 | 25.1761 | 25.1775
| 1:01 | 25.1761 | 25.178
| 1:01 | 25.1764 | 25.1774
| 1:01 | 25.1766 | 25.1773
| 1:01 | 25.1769 | 25.178
| 1:01 | 25.177 | 25.1779
| 1:01 | 25.1769 | 25.1781
| 1:01 | 25.1771 | 25.1786
| 1:01 | 25.1777 | 25.1787
| 1:01 | 25.1777 | 25.179
| 1:01 | 25.1778 | 25.1794
| 1:01 | 25.1783 | 25.1794
| 1:01 | 25.1786 | 25.1795
| 1:01 | 25.179 | 25.18
| 1:01 | 25.1791 | 25.1799
| 1:02 | 25.1788 | 25.1804
| 1:02 | 25.1795 | 25.1808
| 1:02 | 25.1798 | 25.1808
| 1:02 | 25.1797 | 25.1815
| 1:02 | 25.1796 | 25.1813
| 1:02 | 25.1807 | 25.1815
| 1:02 | 25.1804 | 25.1818
| 1:02 | 25.1809 | 25.1822
| 1:02 | 25.1808 | 25.1822
| 1:02 | 25.1812 | 25.1829
| 1:02 | 25.1818 | 25.1831
| 1:02 | 25.1816 | 25.1831
| 1:02 | 25.1816 | 25.1833
| 1:02 | 25.1815 | 25.183
| 1:02 | 25.182 | 25.1827
| 1:02 | 25.1821 | 25.1831
| 1:02 | 25.1816 | 25.1834
| 1:02 | 25.1818 | 25.1833
| 1:02 | 25.1824 | 25.1835
| 1:02 | 25.1827 | 25.1836
| 1:02 | 25.1822 | 25.1838
| 1:02 | 25.1823 | 25.1836
| 1:02 | 25.1826 | 25.1842
| 1:02 | 25.1825 | 25.1845
| 1:02 | 25.1827 | 25.1842
| 1:02 | 25.1829 | 25.1841
| 1:02 | 25.1827 | 25.1837
| 1:02 | 25.1823 | 25.184
| 1:02 | 25.1826 | 25.1841
| 1:02 | 25.1827 | 25.1843
| 1:02 | 25.1825 | 25.1836
| 1:02 | 25.1823 | 25.1836
| 1:02 | 25.1823 | 25.1838
| 1:02 | 25.1819 | 25.1839
| 1:02 | 25.182 | 25.1832
| 1:02 | 25.1818 | 25.1835
| 1:02 | 25.1818 | 25.1835
| 1:02 | 25.1815 | 25.1831
| 1:02 | 25.1823 | 25.1837
| 1:02 | 25.1822 | 25.1841
| 1:02 | 25.182 | 25.1836
| 1:02 | 25.1823 | 25.1836
| 1:02 | 25.1826 | 25.1834
| 1:02 | 25.1824 | 25.1836
| 1:02 | 25.1826 | 25.1837
| 1:02 | 25.1824 | 25.1839
| 1:02 | 25.1823 | 25.1839
| 1:02 | 25.1824 | 25.1838
| 1:02 | 25.1819 | 25.1838
| 1:02 | 25.1822 | 25.184
| 1:02 | 25.1825 | 25.1841
| 1:02 | 25.1824 | 25.1841
| 1:02 | 25.183 | 25.1845
| 1:02 | 25.1825 | 25.1845
| 1:02 | 25.1827 | 25.1848
| 1:02 | 25.1832 | 25.1846
| 1:02 | 25.1832 | 25.1848
| 1:02 | 25.1835 | 25.185
| 1:02 | 25.1832 | 25.1846
| 1:02 | 25.1832 | 25.1846
| 1:03 | 25.1836 | 25.1846
| 1:03 | 25.1833 | 25.1843
| 1:03 | 25.1827 | 25.1851
| 1:03 | 25.183 | 25.1847
| 1:03 | 25.1827 | 25.1844
| 1:03 | 25.1827 | 25.1847
| 1:03 | 25.1829 | 25.1844
| 1:03 | 25.1829 | 25.1844
| 1:03 | 25.1832 | 25.1843
| 1:03 | 25.1831 | 25.1842
| 1:03 | 25.1832 | 25.1843
| 1:03 | 25.1827 | 25.1841
| 1:03 | 25.1832 | 25.1846
| 1:03 | 25.1826 | 25.1838
| 1:03 | 25.1827 | 25.1837
| 1:03 | 25.1827 | 25.1841
| 1:03 | 25.1826 | 25.184
| 1:03 | 25.1821 | 25.184
| 1:03 | 25.1822 | 25.1841
| 1:03 | 25.1826 | 25.1839
| 1:03 | 25.1822 | 25.1838
| 1:03 | 25.1824 | 25.1839
| 1:03 | 25.1825 | 25.1844
| 1:03 | 25.1821 | 25.184
| 1:03 | 25.1823 | 25.1841
| 1:03 | 25.1826 | 25.1843
| 1:03 | 25.1824 | 25.1844
| 1:03 | 25.1824 | 25.1842
| 1:03 | 25.1821 | 25.1839
| 1:03 | 25.1822 | 25.1838
| 1:03 | 25.1819 | 25.1837
| 1:03 | 25.1825 | 25.1836
| 1:03 | 25.1819 | 25.1834
| 1:03 | 25.1821 | 25.1837
| 1:03 | 25.1822 | 25.1841
| 1:03 | 25.1818 | 25.1838
| 1:03 | 25.1824 | 25.1841
| 1:03 | 25.1818 | 25.184
| 1:03 | 25.1823 | 25.1842
| 1:03 | 25.1819 | 25.1836
| 1:03 | 25.1818 | 25.1838
| 1:03 | 25.1825 | 25.1836
| 1:03 | 25.1819 | 25.1833
| 1:03 | 25.1814 | 25.1828
| 1:03 | 25.1815 | 25.183
| 1:03 | 25.1811 | 25.1826
| 1:03 | 25.1811 | 25.1823
| 1:03 | 25.1809 | 25.1824
| 1:03 | 25.1807 | 25.1821
| 1:03 | 25.1801 | 25.1817
| 1:03 | 25.18 | 25.1818
| 1:03 | 25.1799 | 25.1815
| 1:03 | 25.1801 | 25.1813
| 1:03 | 25.1801 | 25.1818
| 1:03 | 25.1795 | 25.1815
| 1:03 | 25.1797 | 25.1819
| 1:03 | 25.1799 | 25.1819
| 1:03 | 25.1805 | 25.1819
| 1:03 | 25.1801 | 25.1821
| 1:03 | 25.1809 | 25.183
| 1:04 | 25.1805 | 25.1826
| 1:04 | 25.1813 | 25.183
| 1:04 | 25.1812 | 25.183
| 1:04 | 25.1816 | 25.1835
| 1:04 | 25.1819 | 25.1835
| 1:04 | 25.1815 | 25.1834
| 1:04 | 25.1816 | 25.1835
| 1:04 | 25.1818 | 25.1838
| 1:04 | 25.1818 | 25.1839
| 1:04 | 25.1821 | 25.1838
| 1:04 | 25.1821 | 25.1844
| 1:04 | 25.1824 | 25.1843
| 1:04 | 25.1826 | 25.1841
| 1:04 | 25.1823 | 25.1846
| 1:04 | 25.1828 | 25.1854
| 1:04 | 25.1826 | 25.1844
| 1:04 | 25.1826 | 25.1848
| 1:04 | 25.1829 | 25.1844
| 1:04 | 25.183 | 25.1853
| 1:04 | 25.1838 | 25.186
| 1:04 | 25.1843 | 25.1867
| 1:04 | 25.1857 | 25.1878
| 1:04 | 25.1864 | 25.1897
| 1:04 | 25.1878 | 25.19
| 1:04 | 25.1883 | 25.1908
| 1:04 | 25.1892 | 25.1915
| 1:04 | 25.1897 | 25.1926
| 1:04 | 25.1905 | 25.1932
| 1:04 | 25.1912 | 25.1937
| 1:04 | 25.1914 | 25.194
| 1:04 | 25.1921 | 25.1942
| 1:04 | 25.1919 | 25.1934
| 1:04 | 25.1914 | 25.1936
| 1:04 | 25.1917 | 25.1932
| 1:04 | 25.191 | 25.1929
| 1:04 | 25.1908 | 25.1925
| 1:04 | 25.1906 | 25.192
| 1:04 | 25.1901 | 25.1917
| 1:04 | 25.1897 | 25.1918
| 1:04 | 25.1895 | 25.1911
| 1:04 | 25.1892 | 25.1906
| 1:04 | 25.1884 | 25.1906
| 1:04 | 25.1882 | 25.1899
| 1:04 | 25.1876 | 25.1893
| 1:04 | 25.1874 | 25.1892
| 1:04 | 25.1872 | 25.1884
| 1:04 | 25.1866 | 25.1877
| 1:04 | 25.1861 | 25.1876
| 1:04 | 25.1857 | 25.187
| 1:04 | 25.1852 | 25.1867
| 1:04 | 25.1849 | 25.1854
| 1:04 | 25.184 | 25.1852
| 1:04 | 25.183 | 25.185
| 1:04 | 25.1829 | 25.1844
| 1:04 | 25.1825 | 25.185
| 1:04 | 25.1828 | 25.184
| 1:04 | 25.1826 | 25.1844
| 1:04 | 25.1827 | 25.1845
| 1:04 | 25.1832 | 25.1845
| 1:04 | 25.1829 | 25.1841
| 1:05 | 25.1828 | 25.1844
| 1:05 | 25.1828 | 25.184
| 1:05 | 25.1823 | 25.1837
| 1:05 | 25.1819 | 25.1832
| 1:05 | 25.1813 | 25.1824
| 1:05 | 25.1812 | 25.1822
| 1:05 | 25.1806 | 25.182
| 1:05 | 25.1795 | 25.181
| 1:05 | 25.1792 | 25.1803
| 1:05 | 25.1783 | 25.1795
| 1:05 | 25.1776 | 25.1788
| 1:05 | 25.1777 | 25.1786
| 1:05 | 25.1771 | 25.1786
| 1:05 | 25.1766 | 25.1781
| 1:05 | 25.1762 | 25.1784
| 1:05 | 25.1762 | 25.178
| 1:05 | 25.1763 | 25.1785
| 1:05 | 25.1766 | 25.1778
| 1:05 | 25.1759 | 25.1779
| 1:05 | 25.1762 | 25.1783
| 1:05 | 25.1762 | 25.1787
| 1:05 | 25.1765 | 25.1783
| 1:05 | 25.1766 | 25.1787
| 1:05 | 25.1766 | 25.1792
| 1:05 | 25.1771 | 25.1793
| 1:05 | 25.1773 | 25.179
| 1:05 | 25.1776 | 25.1794
| 1:05 | 25.1775 | 25.1791
| 1:05 | 25.1772 | 25.1793

## 1e Solution
Description: Your temperature sensor is out of calibration, but you have acquired a temperature sensor that has just recently been calibrated with the below calibration certificate. Find the new expanded measurement uncertainty with a confidence level of 95% for the sensor out of calibration.
Example Application: Temperature sensor reading in °C
Food for Thought: (1) What additional data would you want to collect? (2) What is another component of uncertainty that would be important to consider between calibration or validation cycles?
Answer: (1)Additional data should be taken through the temperature ranges expected. (2) Drift of the calibration over time. Some ways to deal with this is to do a Type A analysis over time or approximate drift with linear degradation from year to year.
Data:
Date | Time | Out-of-Cal Sensor (°C) | Validation Sensor (°C) | Error | Sq. Error
2016-10-14 | 1:00 | 25.1788 | 25.1766 | 0.0022381 | 5.0091e-06 | Validation Sensor Uncertainty | 0.02 | From the Validation Sensor Calibration Certificate: |  | Validation Uncertainty Budget | n | Uncertainty (°C)
| 1:00 | 25.1785 | 25.1769 | 0.0015696 | 2.4635e-06 | Average Error | -0.0011906 |  |  | Validation Sensor Calibration Uncertainty | 100 | 0.02
| 1:00 | 25.1783 | 25.1765 | 0.0017575 | 3.0889e-06 | Validation Repeatability (SSE) | 0.0015228 | Expanded Uncertainty | 0.04 | Validation Repeatability | 300 | 0.0015228
| 1:00 | 25.1787 | 25.1769 | 0.0018332 | 3.3606e-06 | Sensor repeatability | 0.008 | Coverage Factor | 2 | Measurement Uncertainty Budget
| 1:00 | 25.1783 | 25.1767 | 0.0016387 | 2.6853e-06 | Count | 300 | Calibration Uncertainty | 0.02 | Validation Uncertainty | 100 | 0.020058
| 1:00 | 25.1779 | 25.1761 | 0.0018256 | 3.333e-06 |  |  | Expanded Sensor Repeatability | 0.016 | Sensor Repeatability | 100 | 0.008
| 1:00 | 25.1775 | 25.1763 | 0.0012145 | 1.4749e-06 |  |  | Coverage Factor | 2 | Combined Uncertainty |  | 0.021594
| 1:00 | 25.1773 | 25.1758 | 0.0015275 | 2.3331e-06 |  |  | Sensor Repeatability | 0.008 | Degrees of Freedom | 129.72
| 1:00 | 25.1768 | 25.1752 | 0.0016064 | 2.5805e-06 |  |  |  |  | Expansion Factor | 1.9785
| 1:00 | 25.1764 | 25.1749 | 0.0014683 | 2.1559e-06 |  |  |  |  | Expanded Uncertainty |  | 0.042725
| 1:00 | 25.1763 | 25.1752 | 0.0011061 | 1.2235e-06
| 1:00 | 25.1763 | 25.1753 | 0.00098326 | 9.668e-07
| 1:00 | 25.1759 | 25.1749 | 0.00099682 | 9.9366e-07
| 1:00 | 25.176 | 25.1751 | 0.00087732 | 7.6969e-07
| 1:00 | 25.1758 | 25.1748 | 0.0010189 | 1.0381e-06
| 1:00 | 25.1757 | 25.1747 | 0.0010193 | 1.0389e-06
| 1:00 | 25.1756 | 25.1748 | 0.0008186 | 6.7011e-07
| 1:00 | 25.1752 | 25.1752 | 4.1398e-05 | 1.7138e-09
| 1:00 | 25.1757 | 25.1748 | 0.0009209 | 8.4806e-07
| 1:00 | 25.1755 | 25.1746 | 0.00093563 | 8.754e-07
| 1:00 | 25.1753 | 25.1748 | 0.00052675 | 2.7746e-07
| 1:00 | 25.1751 | 25.1741 | 0.00097446 | 9.4957e-07
| 1:00 | 25.1748 | 25.175 | -0.00018303 | 3.3501e-08
| 1:00 | 25.1751 | 25.175 | 0.00014487 | 2.0987e-08
| 1:00 | 25.1754 | 25.1752 | 0.00021042 | 4.4275e-08
| 1:00 | 25.1756 | 25.1752 | 0.00037776 | 1.427e-07
| 1:00 | 25.1754 | 25.1752 | 0.00022953 | 5.2683e-08
| 1:00 | 25.1756 | 25.1751 | 0.00049118 | 2.4126e-07
| 1:00 | 25.1752 | 25.1747 | 0.00051708 | 2.6737e-07
| 1:00 | 25.1749 | 25.175 | -0.00013543 | 1.8341e-08
| 1:01 | 25.1752 | 25.1751 | 0.00012989 | 1.6872e-08
| 1:01 | 25.1752 | 25.1752 | -4.8407e-06 | 2.3432e-11
| 1:01 | 25.175 | 25.1748 | 0.0002021 | 4.0846e-08
| 1:01 | 25.175 | 25.1746 | 0.00043839 | 1.9219e-07
| 1:01 | 25.1747 | 25.1747 | 3.0812e-05 | 9.494e-10
| 1:01 | 25.175 | 25.1746 | 0.00040126 | 1.6101e-07
| 1:01 | 25.1749 | 25.1743 | 0.00058241 | 3.3921e-07
| 1:01 | 25.1745 | 25.1745 | 1.4414e-06 | 2.0776e-12
| 1:01 | 25.1744 | 25.174 | 0.00037989 | 1.4432e-07
| 1:01 | 25.1743 | 25.1742 | 0.00010973 | 1.204e-08
| 1:01 | 25.1741 | 25.1744 | -0.00029543 | 8.7277e-08
| 1:01 | 25.1736 | 25.1741 | -0.00047946 | 2.2988e-07
| 1:01 | 25.1736 | 25.174 | -0.00036446 | 1.3283e-07
| 1:01 | 25.1741 | 25.1745 | -0.00035263 | 1.2435e-07
| 1:01 | 25.1742 | 25.174 | 0.00015708 | 2.4675e-08
| 1:01 | 25.1737 | 25.1743 | -0.0005601 | 3.1371e-07
| 1:01 | 25.1741 | 25.1748 | -0.00072996 | 5.3285e-07
| 1:01 | 25.1745 | 25.1745 | 1.6389e-05 | 2.686e-10
| 1:01 | 25.1742 | 25.1751 | -0.00086889 | 7.5497e-07
| 1:01 | 25.1743 | 25.1755 | -0.0011821 | 1.3975e-06
| 1:01 | 25.1745 | 25.1755 | -0.0010296 | 1.0601e-06
| 1:01 | 25.1749 | 25.1754 | -0.0005342 | 2.8537e-07
| 1:01 | 25.1747 | 25.1752 | -0.00045568 | 2.0764e-07
| 1:01 | 25.1746 | 25.1759 | -0.0013242 | 1.7535e-06
| 1:01 | 25.175 | 25.176 | -0.00096383 | 9.2896e-07
| 1:01 | 25.1751 | 25.1758 | -0.00073928 | 5.4654e-07
| 1:01 | 25.1756 | 25.1761 | -0.00050316 | 2.5317e-07
| 1:01 | 25.1754 | 25.1763 | -0.00090896 | 8.2622e-07
| 1:01 | 25.1756 | 25.1755 | 5.5892e-05 | 3.124e-09
| 1:01 | 25.1755 | 25.176 | -0.00049944 | 2.4944e-07
| 1:01 | 25.1751 | 25.1759 | -0.00080357 | 6.4573e-07
| 1:01 | 25.1752 | 25.1763 | -0.0011429 | 1.3061e-06
| 1:01 | 25.1753 | 25.1762 | -0.00088887 | 7.9008e-07
| 1:01 | 25.1756 | 25.1768 | -0.0011623 | 1.3509e-06
| 1:01 | 25.1759 | 25.1763 | -0.00044026 | 1.9383e-07
| 1:01 | 25.1755 | 25.1766 | -0.0011399 | 1.2994e-06
| 1:01 | 25.1761 | 25.1767 | -0.00056672 | 3.2117e-07
| 1:01 | 25.1763 | 25.1769 | -0.00064418 | 4.1497e-07
| 1:01 | 25.1764 | 25.1774 | -0.0010331 | 1.0673e-06
| 1:01 | 25.1761 | 25.1771 | -0.00096071 | 9.2296e-07
| 1:01 | 25.1764 | 25.1775 | -0.0011068 | 1.225e-06
| 1:01 | 25.1765 | 25.177 | -0.00052849 | 2.793e-07
| 1:01 | 25.177 | 25.1776 | -0.00059158 | 3.4997e-07
| 1:01 | 25.1765 | 25.178 | -0.0015104 | 2.2812e-06
| 1:01 | 25.1767 | 25.1775 | -0.0007863 | 6.1827e-07
| 1:01 | 25.177 | 25.1774 | -0.00042002 | 1.7642e-07
| 1:01 | 25.1761 | 25.1775 | -0.0013955 | 1.9475e-06
| 1:01 | 25.1761 | 25.178 | -0.0018681 | 3.4899e-06
| 1:01 | 25.1764 | 25.1774 | -0.00097593 | 9.5243e-07
| 1:01 | 25.1766 | 25.1773 | -0.00070887 | 5.025e-07
| 1:01 | 25.1769 | 25.178 | -0.0010561 | 1.1154e-06
| 1:01 | 25.177 | 25.1779 | -0.00085391 | 7.2916e-07
| 1:01 | 25.1769 | 25.1781 | -0.0012437 | 1.5468e-06
| 1:01 | 25.1771 | 25.1786 | -0.0015081 | 2.2745e-06
| 1:01 | 25.1777 | 25.1787 | -0.00095932 | 9.2029e-07
| 1:01 | 25.1777 | 25.179 | -0.0013052 | 1.7036e-06
| 1:01 | 25.1778 | 25.1794 | -0.001552 | 2.4088e-06
| 1:01 | 25.1783 | 25.1794 | -0.0011498 | 1.3219e-06
| 1:01 | 25.1786 | 25.1795 | -0.00089462 | 8.0034e-07
| 1:01 | 25.179 | 25.18 | -0.0010116 | 1.0233e-06
| 1:01 | 25.1791 | 25.1799 | -0.00084641 | 7.1641e-07
| 1:02 | 25.1788 | 25.1804 | -0.0015734 | 2.4755e-06
| 1:02 | 25.1795 | 25.1808 | -0.0013241 | 1.7532e-06
| 1:02 | 25.1798 | 25.1808 | -0.0010011 | 1.0023e-06
| 1:02 | 25.1797 | 25.1815 | -0.0018432 | 3.3972e-06
| 1:02 | 25.1796 | 25.1813 | -0.0016569 | 2.7454e-06
| 1:02 | 25.1807 | 25.1815 | -0.00084409 | 7.1248e-07
| 1:02 | 25.1804 | 25.1818 | -0.001412 | 1.9937e-06
| 1:02 | 25.1809 | 25.1822 | -0.0013151 | 1.7294e-06
| 1:02 | 25.1808 | 25.1822 | -0.0014276 | 2.0379e-06
| 1:02 | 25.1812 | 25.1829 | -0.0017328 | 3.0024e-06
| 1:02 | 25.1818 | 25.1831 | -0.0012568 | 1.5796e-06
| 1:02 | 25.1816 | 25.1831 | -0.0015294 | 2.3392e-06
| 1:02 | 25.1816 | 25.1833 | -0.0016958 | 2.8756e-06
| 1:02 | 25.1815 | 25.183 | -0.0014813 | 2.1942e-06
| 1:02 | 25.182 | 25.1827 | -0.00073792 | 5.4453e-07
| 1:02 | 25.1821 | 25.1831 | -0.0010094 | 1.0189e-06
| 1:02 | 25.1816 | 25.1834 | -0.0017517 | 3.0685e-06
| 1:02 | 25.1818 | 25.1833 | -0.0014689 | 2.1578e-06
| 1:02 | 25.1824 | 25.1835 | -0.0011303 | 1.2777e-06
| 1:02 | 25.1827 | 25.1836 | -0.00086642 | 7.5069e-07
| 1:02 | 25.1822 | 25.1838 | -0.0016366 | 2.6786e-06
| 1:02 | 25.1823 | 25.1836 | -0.0013259 | 1.7581e-06
| 1:02 | 25.1826 | 25.1842 | -0.0015981 | 2.5539e-06
| 1:02 | 25.1825 | 25.1845 | -0.002016 | 4.0641e-06
| 1:02 | 25.1827 | 25.1842 | -0.0014923 | 2.2268e-06
| 1:02 | 25.1829 | 25.1841 | -0.0012136 | 1.4729e-06
| 1:02 | 25.1827 | 25.1837 | -0.00098665 | 9.7348e-07
| 1:02 | 25.1823 | 25.184 | -0.0017273 | 2.9836e-06
| 1:02 | 25.1826 | 25.1841 | -0.0015092 | 2.2776e-06
| 1:02 | 25.1827 | 25.1843 | -0.0016247 | 2.6395e-06
| 1:02 | 25.1825 | 25.1836 | -0.0010501 | 1.1027e-06
| 1:02 | 25.1823 | 25.1836 | -0.0012731 | 1.6207e-06
| 1:02 | 25.1823 | 25.1838 | -0.0014585 | 2.1272e-06
| 1:02 | 25.1819 | 25.1839 | -0.0019516 | 3.8087e-06
| 1:02 | 25.182 | 25.1832 | -0.0012184 | 1.4845e-06
| 1:02 | 25.1818 | 25.1835 | -0.0016641 | 2.7691e-06
| 1:02 | 25.1818 | 25.1835 | -0.0017449 | 3.0445e-06
| 1:02 | 25.1815 | 25.1831 | -0.0016017 | 2.5654e-06
| 1:02 | 25.1823 | 25.1837 | -0.0014499 | 2.1021e-06
| 1:02 | 25.1822 | 25.1841 | -0.0018751 | 3.516e-06
| 1:02 | 25.182 | 25.1836 | -0.0016079 | 2.5854e-06
| 1:02 | 25.1823 | 25.1836 | -0.0012831 | 1.6463e-06
| 1:02 | 25.1826 | 25.1834 | -0.000826 | 6.8228e-07
| 1:02 | 25.1824 | 25.1836 | -0.0011692 | 1.3671e-06
| 1:02 | 25.1826 | 25.1837 | -0.0010988 | 1.2074e-06
| 1:02 | 25.1824 | 25.1839 | -0.0015154 | 2.2965e-06
| 1:02 | 25.1823 | 25.1839 | -0.001616 | 2.6116e-06
| 1:02 | 25.1824 | 25.1838 | -0.0013674 | 1.8699e-06
| 1:02 | 25.1819 | 25.1838 | -0.0019463 | 3.7882e-06
| 1:02 | 25.1822 | 25.184 | -0.0018245 | 3.329e-06
| 1:02 | 25.1825 | 25.1841 | -0.0016079 | 2.5854e-06
| 1:02 | 25.1824 | 25.1841 | -0.0017415 | 3.033e-06
| 1:02 | 25.183 | 25.1845 | -0.0015067 | 2.2701e-06
| 1:02 | 25.1825 | 25.1845 | -0.0020429 | 4.1736e-06
| 1:02 | 25.1827 | 25.1848 | -0.0021347 | 4.557e-06
| 1:02 | 25.1832 | 25.1846 | -0.0014017 | 1.9646e-06
| 1:02 | 25.1832 | 25.1848 | -0.0016022 | 2.567e-06
| 1:02 | 25.1835 | 25.185 | -0.001462 | 2.1375e-06
| 1:02 | 25.1832 | 25.1846 | -0.0013577 | 1.8434e-06
| 1:02 | 25.1832 | 25.1846 | -0.0014166 | 2.0068e-06
| 1:03 | 25.1836 | 25.1846 | -0.00098761 | 9.7538e-07
| 1:03 | 25.1833 | 25.1843 | -0.00099968 | 9.9936e-07
| 1:03 | 25.1827 | 25.1851 | -0.0023676 | 5.6056e-06
| 1:03 | 25.183 | 25.1847 | -0.0017069 | 2.9136e-06
| 1:03 | 25.1827 | 25.1844 | -0.0017076 | 2.9158e-06
| 1:03 | 25.1827 | 25.1847 | -0.0020441 | 4.1782e-06
| 1:03 | 25.1829 | 25.1844 | -0.0014687 | 2.1572e-06
| 1:03 | 25.1829 | 25.1844 | -0.0015082 | 2.2746e-06
| 1:03 | 25.1832 | 25.1843 | -0.0010926 | 1.1938e-06
| 1:03 | 25.1831 | 25.1842 | -0.0011007 | 1.2116e-06
| 1:03 | 25.1832 | 25.1843 | -0.0011257 | 1.2673e-06
| 1:03 | 25.1827 | 25.1841 | -0.001356 | 1.8387e-06
| 1:03 | 25.1832 | 25.1846 | -0.0013901 | 1.9324e-06
| 1:03 | 25.1826 | 25.1838 | -0.0011964 | 1.4314e-06
| 1:03 | 25.1827 | 25.1837 | -0.0010465 | 1.0951e-06
| 1:03 | 25.1827 | 25.1841 | -0.0014048 | 1.9736e-06
| 1:03 | 25.1826 | 25.184 | -0.0014201 | 2.0166e-06
| 1:03 | 25.1821 | 25.184 | -0.0018966 | 3.5973e-06
| 1:03 | 25.1822 | 25.1841 | -0.0018685 | 3.4911e-06
| 1:03 | 25.1826 | 25.1839 | -0.0013239 | 1.7527e-06
| 1:03 | 25.1822 | 25.1838 | -0.0015514 | 2.4069e-06
| 1:03 | 25.1824 | 25.1839 | -0.0014522 | 2.1089e-06
| 1:03 | 25.1825 | 25.1844 | -0.0018542 | 3.4382e-06
| 1:03 | 25.1821 | 25.184 | -0.0018624 | 3.4686e-06
| 1:03 | 25.1823 | 25.1841 | -0.0017982 | 3.2336e-06
| 1:03 | 25.1826 | 25.1843 | -0.0017433 | 3.0391e-06
| 1:03 | 25.1824 | 25.1844 | -0.0020057 | 4.023e-06
| 1:03 | 25.1824 | 25.1842 | -0.0018362 | 3.3715e-06
| 1:03 | 25.1821 | 25.1839 | -0.0018479 | 3.4148e-06
| 1:03 | 25.1822 | 25.1838 | -0.0016296 | 2.6555e-06
| 1:03 | 25.1819 | 25.1837 | -0.0017894 | 3.2021e-06
| 1:03 | 25.1825 | 25.1836 | -0.0011105 | 1.2332e-06
| 1:03 | 25.1819 | 25.1834 | -0.0015114 | 2.2842e-06
| 1:03 | 25.1821 | 25.1837 | -0.00163 | 2.6568e-06
| 1:03 | 25.1822 | 25.1841 | -0.0018686 | 3.4917e-06
| 1:03 | 25.1818 | 25.1838 | -0.001999 | 3.9962e-06
| 1:03 | 25.1824 | 25.1841 | -0.0017004 | 2.8913e-06
| 1:03 | 25.1818 | 25.184 | -0.0021794 | 4.7496e-06
| 1:03 | 25.1823 | 25.1842 | -0.0018525 | 3.4317e-06
| 1:03 | 25.1819 | 25.1836 | -0.0016527 | 2.7315e-06
| 1:03 | 25.1818 | 25.1838 | -0.0020257 | 4.1035e-06
| 1:03 | 25.1825 | 25.1836 | -0.0011327 | 1.283e-06
| 1:03 | 25.1819 | 25.1833 | -0.0013765 | 1.8947e-06
| 1:03 | 25.1814 | 25.1828 | -0.0014375 | 2.0664e-06
| 1:03 | 25.1815 | 25.183 | -0.0015101 | 2.2805e-06
| 1:03 | 25.1811 | 25.1826 | -0.0014956 | 2.2369e-06
| 1:03 | 25.1811 | 25.1823 | -0.0012327 | 1.5195e-06
| 1:03 | 25.1809 | 25.1824 | -0.001488 | 2.2141e-06
| 1:03 | 25.1807 | 25.1821 | -0.0014101 | 1.9884e-06
| 1:03 | 25.1801 | 25.1817 | -0.0016284 | 2.6518e-06
| 1:03 | 25.18 | 25.1818 | -0.0017958 | 3.2249e-06
| 1:03 | 25.1799 | 25.1815 | -0.0015983 | 2.5546e-06
| 1:03 | 25.1801 | 25.1813 | -0.0012261 | 1.5034e-06
| 1:03 | 25.1801 | 25.1818 | -0.0016872 | 2.8466e-06
| 1:03 | 25.1795 | 25.1815 | -0.0019672 | 3.8698e-06
| 1:03 | 25.1797 | 25.1819 | -0.0021846 | 4.7725e-06
| 1:03 | 25.1799 | 25.1819 | -0.0020389 | 4.1569e-06
| 1:03 | 25.1805 | 25.1819 | -0.0013752 | 1.8911e-06
| 1:03 | 25.1801 | 25.1821 | -0.0020217 | 4.0871e-06
| 1:03 | 25.1809 | 25.183 | -0.0020677 | 4.2756e-06
| 1:04 | 25.1805 | 25.1826 | -0.0021023 | 4.4196e-06
| 1:04 | 25.1813 | 25.183 | -0.0017043 | 2.9045e-06
| 1:04 | 25.1812 | 25.183 | -0.0018211 | 3.3163e-06
| 1:04 | 25.1816 | 25.1835 | -0.0019005 | 3.6117e-06
| 1:04 | 25.1819 | 25.1835 | -0.0016364 | 2.6777e-06
| 1:04 | 25.1815 | 25.1834 | -0.0019217 | 3.6929e-06
| 1:04 | 25.1816 | 25.1835 | -0.001892 | 3.5796e-06
| 1:04 | 25.1818 | 25.1838 | -0.0019808 | 3.9237e-06
| 1:04 | 25.1818 | 25.1839 | -0.00205 | 4.2027e-06
| 1:04 | 25.1821 | 25.1838 | -0.001679 | 2.819e-06
| 1:04 | 25.1821 | 25.1844 | -0.0022504 | 5.0643e-06
| 1:04 | 25.1824 | 25.1843 | -0.0018724 | 3.5059e-06
| 1:04 | 25.1826 | 25.1841 | -0.0015158 | 2.2976e-06
| 1:04 | 25.1823 | 25.1846 | -0.0022779 | 5.1889e-06
| 1:04 | 25.1828 | 25.1854 | -0.0026456 | 6.9991e-06
| 1:04 | 25.1826 | 25.1844 | -0.0017945 | 3.2202e-06
| 1:04 | 25.1826 | 25.1848 | -0.0021613 | 4.6714e-06
| 1:04 | 25.1829 | 25.1844 | -0.00149 | 2.2201e-06
| 1:04 | 25.183 | 25.1853 | -0.0023133 | 5.3515e-06
| 1:04 | 25.1838 | 25.186 | -0.0022421 | 5.027e-06
| 1:04 | 25.1843 | 25.1867 | -0.0023778 | 5.6539e-06
| 1:04 | 25.1857 | 25.1878 | -0.0021174 | 4.4833e-06
| 1:04 | 25.1864 | 25.1897 | -0.003336 | 1.1129e-05
| 1:04 | 25.1878 | 25.19 | -0.0021562 | 4.6494e-06
| 1:04 | 25.1883 | 25.1908 | -0.0025242 | 6.3718e-06
| 1:04 | 25.1892 | 25.1915 | -0.0022627 | 5.1198e-06
| 1:04 | 25.1897 | 25.1926 | -0.0028793 | 8.2904e-06
| 1:04 | 25.1905 | 25.1932 | -0.00275 | 7.5624e-06
| 1:04 | 25.1912 | 25.1937 | -0.0025409 | 6.4563e-06
| 1:04 | 25.1914 | 25.194 | -0.0025774 | 6.6428e-06
| 1:04 | 25.1921 | 25.1942 | -0.0021192 | 4.4911e-06
| 1:04 | 25.1919 | 25.1934 | -0.0014711 | 2.1642e-06
| 1:04 | 25.1914 | 25.1936 | -0.0022062 | 4.8673e-06
| 1:04 | 25.1917 | 25.1932 | -0.0015014 | 2.2543e-06
| 1:04 | 25.191 | 25.1929 | -0.0019139 | 3.6632e-06
| 1:04 | 25.1908 | 25.1925 | -0.0017394 | 3.0254e-06
| 1:04 | 25.1906 | 25.192 | -0.0013883 | 1.9274e-06
| 1:04 | 25.1901 | 25.1917 | -0.0016328 | 2.6659e-06
| 1:04 | 25.1897 | 25.1918 | -0.00213 | 4.537e-06
| 1:04 | 25.1895 | 25.1911 | -0.0016045 | 2.5743e-06
| 1:04 | 25.1892 | 25.1906 | -0.0013714 | 1.8809e-06
| 1:04 | 25.1884 | 25.1906 | -0.0022265 | 4.9571e-06
| 1:04 | 25.1882 | 25.1899 | -0.0016722 | 2.7964e-06
| 1:04 | 25.1876 | 25.1893 | -0.0017357 | 3.0126e-06
| 1:04 | 25.1874 | 25.1892 | -0.0018279 | 3.3413e-06
| 1:04 | 25.1872 | 25.1884 | -0.001159 | 1.3433e-06
| 1:04 | 25.1866 | 25.1877 | -0.0011408 | 1.3014e-06
| 1:04 | 25.1861 | 25.1876 | -0.0015265 | 2.3303e-06
| 1:04 | 25.1857 | 25.187 | -0.0012526 | 1.5691e-06
| 1:04 | 25.1852 | 25.1867 | -0.001454 | 2.1143e-06
| 1:04 | 25.1849 | 25.1854 | -0.00052338 | 2.7393e-07
| 1:04 | 25.184 | 25.1852 | -0.0012457 | 1.5517e-06
| 1:04 | 25.183 | 25.185 | -0.0019505 | 3.8043e-06
| 1:04 | 25.1829 | 25.1844 | -0.0014916 | 2.2248e-06
| 1:04 | 25.1825 | 25.185 | -0.0024856 | 6.1783e-06
| 1:04 | 25.1828 | 25.184 | -0.0012028 | 1.4467e-06
| 1:04 | 25.1826 | 25.1844 | -0.0018366 | 3.373e-06
| 1:04 | 25.1827 | 25.1845 | -0.0017614 | 3.1025e-06
| 1:04 | 25.1832 | 25.1845 | -0.0012967 | 1.6814e-06
| 1:04 | 25.1829 | 25.1841 | -0.0012317 | 1.517e-06
| 1:05 | 25.1828 | 25.1844 | -0.0016181 | 2.6182e-06
| 1:05 | 25.1828 | 25.184 | -0.0012238 | 1.4976e-06
| 1:05 | 25.1823 | 25.1837 | -0.0013751 | 1.8909e-06
| 1:05 | 25.1819 | 25.1832 | -0.0013255 | 1.7569e-06
| 1:05 | 25.1813 | 25.1824 | -0.0010777 | 1.1614e-06
| 1:05 | 25.1812 | 25.1822 | -0.0010306 | 1.0622e-06
| 1:05 | 25.1806 | 25.182 | -0.001401 | 1.9627e-06
| 1:05 | 25.1795 | 25.181 | -0.0015479 | 2.396e-06
| 1:05 | 25.1792 | 25.1803 | -0.0010796 | 1.1656e-06
| 1:05 | 25.1783 | 25.1795 | -0.0012188 | 1.4855e-06
| 1:05 | 25.1776 | 25.1788 | -0.0011721 | 1.3739e-06
| 1:05 | 25.1777 | 25.1786 | -0.00090351 | 8.1633e-07
| 1:05 | 25.1771 | 25.1786 | -0.0014739 | 2.1725e-06
| 1:05 | 25.1766 | 25.1781 | -0.0015221 | 2.3169e-06
| 1:05 | 25.1762 | 25.1784 | -0.0021604 | 4.6675e-06
| 1:05 | 25.1762 | 25.178 | -0.0018025 | 3.249e-06
| 1:05 | 25.1763 | 25.1785 | -0.0021996 | 4.8384e-06
| 1:05 | 25.1766 | 25.1778 | -0.0012434 | 1.5461e-06
| 1:05 | 25.1759 | 25.1779 | -0.0020088 | 4.0352e-06
| 1:05 | 25.1762 | 25.1783 | -0.0020501 | 4.2028e-06
| 1:05 | 25.1762 | 25.1787 | -0.002523 | 6.3655e-06
| 1:05 | 25.1765 | 25.1783 | -0.0018416 | 3.3916e-06
| 1:05 | 25.1766 | 25.1787 | -0.0021446 | 4.5991e-06
| 1:05 | 25.1766 | 25.1792 | -0.0025657 | 6.5827e-06
| 1:05 | 25.1771 | 25.1793 | -0.0021746 | 4.7291e-06
| 1:05 | 25.1773 | 25.179 | -0.0016845 | 2.8377e-06
| 1:05 | 25.1776 | 25.1794 | -0.0018471 | 3.4119e-06
| 1:05 | 25.1775 | 25.1791 | -0.0016426 | 2.698e-06
| 1:05 | 25.1772 | 25.1793 | -0.0020588 | 4.2385e-06

## Exercise 2a
Description: Blank filters are weighed on a scale. The scientist wants to use an average of the last 80 filters as the blank reading for all filters of the same type (same manufacturer and part number). The scale's calibration certificate is provided below. What is the expanded measurement uncertainty with a confidence level of 95% for using an average blank value?
Example Application: PM10 particulate collection filters.
Food for Thought: (1) Do you use standard deviation or standard error for this analysis? (2) What additional information might you need to ensure this is a proper analysis?
Data:
Date | Filter Pre-weight (mg)
2015-03-01 | 135.45
2015-03-07 | 136.3049
2015-03-13 | 135.7995
2015-03-19 | 135.4888
2015-03-25 | 135.3573
2015-03-31 | 135.4784
2015-04-06 | 135.5135
2015-04-12 | 135.5215
2015-04-18 | 135.1755
2015-04-24 | 135.0218
2015-04-30 | 135.2287
2015-05-06 | 135.4849
2015-05-12 | 135.792
2015-05-18 | 134.9838
2015-05-24 | 135.5607
2015-05-30 | 135.65
2015-06-05 | 135.8347
2015-06-11 | 135.5308
2015-06-17 | 135.5665
2015-06-23 | 135.3603
2015-06-29 | 136.1732
2015-07-05 | 135.6407
2015-07-11 | 135.6118
2015-07-17 | 135.8227
2015-07-23 | 135.4522
2015-07-29 | 135.3028
2015-08-04 | 135.3277
2015-08-10 | 135.7973
2015-08-16 | 135.6123
2015-08-22 | 135.3364
2015-08-28 | 135.2869
2015-09-03 | 135.8288
2015-09-09 | 135.4157
2015-09-15 | 135.8036
2015-09-21 | 135.7351
2015-09-27 | 135.4009
2015-10-03 | 135.6633
2015-10-09 | 135.1611
2015-10-15 | 135.6099
2015-10-21 | 135.4695
2015-10-27 | 135.1065
2015-11-02 | 135.7834
2015-11-08 | 135.3693
2015-11-14 | 135.7481
2015-11-20 | 135.1147
2015-11-26 | 135.5721
2015-12-02 | 135.5737
2015-12-08 | 135.4232
2015-12-14 | 135.5323
2015-12-20 | 135.848
2015-12-26 | 135.3074
2016-01-01 | 135.5799
2016-01-07 | 135.338
2016-01-13 | 135.9204
2016-01-19 | 135.3368
2016-01-25 | 135.5622
2016-01-31 | 135.6835
2016-02-06 | 135.4802
2016-02-12 | 135.8103
2016-02-18 | 135.7168
2016-02-24 | 135.558
2016-03-01 | 135.3512
2016-03-07 | 135.449
2016-03-13 | 135.483
2016-03-19 | 135.3704
2016-03-25 | 135.9062
2016-03-31 | 135.2178
2016-04-06 | 135.3024
2016-04-12 | 135.381
2016-04-18 | 135.4394
2016-04-24 | 135.7064
2016-04-30 | 135.8525
2016-05-06 | 135.6559
2016-05-12 | 135.1548
2016-05-18 | 135.3512
2016-05-24 | 135.8562
2016-05-30 | 135.4316
2016-06-05 | 135.8512
2016-06-11 | 135.6148
2016-06-17 | 135.4016

## 2a Solution
Description: Blank filters are weighed on a scale. The scientist wants to use an average of the last 80 filters as the blank reading for all filters of the same type (same manufacturer and part number). The scale's calibration certificate is provided below. What is the expanded measurement uncertainty with a confidence level of 95% for using an average blank value?
Example Application: PM10 particulate collection filters.
Food for Thought: (1) Do you use standard deviation or standard error for this analysis? (2) What additional information might you need to ensure this is a proper analysis?
Answer: (1) You use standard deviation because one filter will be used, not a collection of filters. (2) You don't know the quantity of filters per lot. If the manufacturer provided 100 filters/lot and the data only spans 2 lots, you may not be properly accounting for the variability between lots. If the filters come in as 10 filters per lot, this method would be okay if the distribution is normal.
From the Scale Calibration Certificate:
Data: |  | Analysis from the Data:
Date | Filter Pre-weight (mg) |  |  | Expanded Uncertainty | 0.0058
2015-03-01 | 135.45 | average | 135.5337 | Coverage Factor | 2
2015-03-07 | 136.3049 | standard deviation | 0.24819 | Calibration Uncertainty | 0.0029
2015-03-13 | 135.7995 | count | 80
2015-03-19 | 135.4888
2015-03-25 | 135.3573 |  |  | Uncertainty Components | n | Uncertainty (mg)
2015-03-31 | 135.4784 |  |  | Calibration Uncertainty | 100 | 0.0029
2015-04-06 | 135.5135 |  |  | Repeatability | 80 | 0.24819
2015-04-12 | 135.5215 | Bin | Frequency | Combined Uncertainty |  | 0.24821
2015-04-18 | 135.1755 | 134.9838 | 1 | Degrees of Freedom | 79.0216
2015-04-24 | 135.0218 | 135.1489 | 3 | Expansion Factor | 1.9905
2015-04-30 | 135.2287 | 135.3141 | 9 | Expanded Uncertainty |  | 0.49404
2015-05-06 | 135.4849 | 135.4792 | 22
2015-05-12 | 135.792 | 135.6444 | 20
2015-05-18 | 134.9838 | 135.8095 | 13
2015-05-24 | 135.5607 | 135.9746 | 10
2015-05-30 | 135.65 | 136.1398 | 0
2015-06-05 | 135.8347 | More | 2
2015-06-11 | 135.5308
2015-06-17 | 135.5665
2015-06-23 | 135.3603
2015-06-29 | 136.1732
2015-07-05 | 135.6407
2015-07-11 | 135.6118
2015-07-17 | 135.8227
2015-07-23 | 135.4522
2015-07-29 | 135.3028
2015-08-04 | 135.3277
2015-08-10 | 135.7973
2015-08-16 | 135.6123
2015-08-22 | 135.3364
2015-08-28 | 135.2869
2015-09-03 | 135.8288
2015-09-09 | 135.4157
2015-09-15 | 135.8036
2015-09-21 | 135.7351
2015-09-27 | 135.4009
2015-10-03 | 135.6633
2015-10-09 | 135.1611
2015-10-15 | 135.6099
2015-10-21 | 135.4695
2015-10-27 | 135.1065
2015-11-02 | 135.7834
2015-11-08 | 135.3693
2015-11-14 | 135.7481
2015-11-20 | 135.1147
2015-11-26 | 135.5721
2015-12-02 | 135.5737
2015-12-08 | 135.4232
2015-12-14 | 135.5323
2015-12-20 | 135.848
2015-12-26 | 135.3074
2016-01-01 | 135.5799
2016-01-07 | 135.338
2016-01-13 | 135.9204
2016-01-19 | 135.3368
2016-01-25 | 135.5622
2016-01-31 | 135.6835
2016-02-06 | 135.4802
2016-02-12 | 135.8103
2016-02-18 | 135.7168
2016-02-24 | 135.558
2016-03-01 | 135.3512
2016-03-07 | 135.449
2016-03-13 | 135.483
2016-03-19 | 135.3704
2016-03-25 | 135.9062
2016-03-31 | 135.2178
2016-04-06 | 135.3024
2016-04-12 | 135.381
2016-04-18 | 135.4394
2016-04-24 | 135.7064
2016-04-30 | 135.8525
2016-05-06 | 135.6559
2016-05-12 | 135.1548
2016-05-18 | 135.3512
2016-05-24 | 135.8562
2016-05-30 | 135.4316
2016-06-05 | 135.8512
2016-06-11 | 135.6148
2016-06-17 | 135.4016

## Exercise 2b
Description: Blank filters are weighed on a scale. The scientist wants to use an average of the last 80 filters as the blank reading for all filters of the same type (same manufacturer and part number). The scale's calibration certificate is provided below. What is the estimated expanded RELATIVE uncertainty with a 95% confidence level for using an average blank value?
Example Application: PM10 particulate collection filters.
Food for Thought: (1) What is wrong with this uncertainty analysis? (2) How could you compute uncertainty for the calibration if the expanded uncertainty was not provided?
Data:
Date | Filter Pre-weight (mg)
2015-03-01 | 135.45
2015-03-07 | 136.3049
2015-03-13 | 135.7995
2015-03-19 | 135.4888
2015-03-25 | 135.3573
2015-03-31 | 135.4784
2015-04-06 | 135.5135
2015-04-12 | 135.5215
2015-04-18 | 135.1755
2015-04-24 | 135.0218
2015-04-30 | 135.2287
2015-05-06 | 135.4849
2015-05-12 | 135.792
2015-05-18 | 134.9838
2015-05-24 | 135.5607
2015-05-30 | 135.65
2015-06-05 | 135.8347
2015-06-11 | 135.5308
2015-06-17 | 135.5665
2015-06-23 | 135.3603
2015-06-29 | 136.1732
2015-07-05 | 135.6407
2015-07-11 | 135.6118
2015-07-17 | 135.8227
2015-07-23 | 135.4522
2015-07-29 | 135.3028
2015-08-04 | 135.3277
2015-08-10 | 135.7973
2015-08-16 | 135.6123
2015-08-22 | 135.3364
2015-08-28 | 135.2869
2015-09-03 | 135.8288
2015-09-09 | 135.4157
2015-09-15 | 135.8036
2015-09-21 | 135.7351
2015-09-27 | 135.4009
2015-10-03 | 135.6633
2015-10-09 | 135.1611
2015-10-15 | 135.6099
2015-10-21 | 135.4695
2015-10-27 | 135.1065
2015-11-02 | 135.7834
2015-11-08 | 135.3693
2015-11-14 | 135.7481
2015-11-20 | 135.1147
2015-11-26 | 135.5721
2015-12-02 | 135.5737
2015-12-08 | 135.4232
2015-12-14 | 135.5323
2015-12-20 | 135.848
2015-12-26 | 135.3074
2016-01-01 | 135.5799
2016-01-07 | 135.338
2016-01-13 | 135.9204
2016-01-19 | 135.3368
2016-01-25 | 135.5622
2016-01-31 | 135.6835
2016-02-06 | 135.4802
2016-02-12 | 135.8103
2016-02-18 | 135.7168
2016-02-24 | 135.558
2016-03-01 | 135.3512
2016-03-07 | 135.449
2016-03-13 | 135.483
2016-03-19 | 135.3704
2016-03-25 | 135.9062
2016-03-31 | 135.2178
2016-04-06 | 135.3024
2016-04-12 | 135.381
2016-04-18 | 135.4394
2016-04-24 | 135.7064
2016-04-30 | 135.8525
2016-05-06 | 135.6559
2016-05-12 | 135.1548
2016-05-18 | 135.3512
2016-05-24 | 135.8562
2016-05-30 | 135.4316
2016-06-05 | 135.8512
2016-06-11 | 135.6148
2016-06-17 | 135.4016

## 2b Solution
Description: Blank filters are weighed on a scale. The scientist wants to use an average of the last 80 filters as the blank reading for all filters of the same type (same manufacturer and part number). The scale's calibration certificate is provided below. What is the estimated expanded RELATIVE uncertainty with a 95% confidence level for using an average blank value?
Example Application: PM10 particulate collection filters.
Food for Thought: (1) What is wrong with this uncertainty analysis? (2) How could you compute uncertainty for the calibration if the expanded uncertainty was not provided?
Answer: (1) The calibration of the scale is expired during some of the analysis. The uncertainty estimated by the manufacturer is no longer guaranteed, therefore only include the dates that are within the valid date range. (2) Use the tolerance and covert it to uncertainty by deviding it be the square root of 3 (See Equation Sheet).
|  |  |  | Using Uncertainty |  |  | (2) If only tolerance was provided
Data: |  | Analysis from the Data: |  | From the Scale Calibration Certificate: |  |  | From the Scale Calibration Certificate:
Date | Filter Pre-weight (mg)
2015-03-01 | 135.45 | average | 135.5364 | Expanded Uncertainty | 0.0058 |  | Tolerance | 0.003
2015-03-07 | 136.3049 | standard deviation | 0.25848 | Coverage Factor | 2 |  | Weight | 100
2015-03-13 | 135.7995 | count | 59 | Calibration Uncertainty | 0.0029 |  | Calibration Uncertainty | 0.0017321
2015-03-19 | 135.4888
2015-03-25 | 135.3573 | [7 empty] | Uncertainty (%)
2015-03-31 | 135.4784 |  |  | Uncertainty Components | n | Uncertainty (%) | Uncertainty Components | n | 0.001278%
2015-04-06 | 135.5135 |  |  | Calibration Uncertainty | 100 | 0.00214% | Calibration Uncertainty | 100 | 0.1907%
2015-04-12 | 135.5215 | Bin | Frequency | Repeatability | 59 | 0.1907% | Repeatability | 59 | 0.1907%
2015-04-18 | 135.1755 | 134.9838 | 1 | Combined Uncertainty |  | 0.1907% | Combined Uncertainty
2015-04-24 | 135.0218 | 135.1725 | 4 | Degrees of Freedom | 58.0146 |  | Degrees of Freedom | 58.0052
2015-04-30 | 135.2287 | 135.3613 | 11 | Expansion Factor | 2.0017 |  | Expansion Factor | 2.0017 | 0.3818%
2015-05-06 | 135.4849 | 135.55 | 15 | Expanded Uncertainty |  | 0.3818% | Expanded Uncertainty
2015-05-12 | 135.792 | 135.7387 | 14
2015-05-18 | 134.9838 | 135.9274 | 12
2015-05-24 | 135.5607 | 136.1162 | 0
2015-05-30 | 135.65 | More | 2
2015-06-05 | 135.8347
2015-06-11 | 135.5308
2015-06-17 | 135.5665
2015-06-23 | 135.3603
2015-06-29 | 136.1732
2015-07-05 | 135.6407
2015-07-11 | 135.6118
2015-07-17 | 135.8227
2015-07-23 | 135.4522
2015-07-29 | 135.3028
2015-08-04 | 135.3277
2015-08-10 | 135.7973
2015-08-16 | 135.6123
2015-08-22 | 135.3364
2015-08-28 | 135.2869
2015-09-03 | 135.8288
2015-09-09 | 135.4157
2015-09-15 | 135.8036
2015-09-21 | 135.7351
2015-09-27 | 135.4009
2015-10-03 | 135.6633
2015-10-09 | 135.1611
2015-10-15 | 135.6099
2015-10-21 | 135.4695
2015-10-27 | 135.1065
2015-11-02 | 135.7834
2015-11-08 | 135.3693
2015-11-14 | 135.7481
2015-11-20 | 135.1147
2015-11-26 | 135.5721
2015-12-02 | 135.5737
2015-12-08 | 135.4232
2015-12-14 | 135.5323
2015-12-20 | 135.848
2015-12-26 | 135.3074
2016-01-01 | 135.5799
2016-01-07 | 135.338
2016-01-13 | 135.9204
2016-01-19 | 135.3368
2016-01-25 | 135.5622
2016-01-31 | 135.6835
2016-02-06 | 135.4802
2016-02-12 | 135.8103
2016-02-18 | 135.7168
2016-02-24 | 135.558
2016-03-01 | 135.3512
2016-03-07 | 135.449
2016-03-13 | 135.483
2016-03-19 | 135.3704
2016-03-25 | 135.9062
2016-03-31 | 135.2178
2016-04-06 | 135.3024
2016-04-12 | 135.381
2016-04-18 | 135.4394
2016-04-24 | 135.7064
2016-04-30 | 135.8525
2016-05-06 | 135.6559
2016-05-12 | 135.1548
2016-05-18 | 135.3512
2016-05-24 | 135.8562
2016-05-30 | 135.4316
2016-06-05 | 135.8512
2016-06-11 | 135.6148
2016-06-17 | 135.4016

## Exercise 2c
Description: Blank filters are weighed on a scale. The scientist wants to use an average of the last 80 filters as the blank reading for all filters of the same type (same manufacturer and part number). The scientist knows there was a new lot opened when they initiated research on March 1, 2015 and there are 10 filters per lot. The scale's calibration certificate is provided below. What is the expanded measurement uncertainty with a confidence level of 95% for using an average blank value?
Example Application: PM10 particulate collection filters.
Food for Thought: (1) What is wrong with this uncertainty analysis? Hint: Check the distribution. (2) Can the scientist use the average approach in this case?
Data:
Date | Filter Pre-weight (mg)
2015-03-01 | 135.1632 | *New lot
2015-03-07 | 135.0692
2015-03-13 | 135.1123
2015-03-19 | 135.2624
2015-03-25 | 135.0345
2015-03-31 | 135.1475
2015-04-06 | 134.9215
2015-04-12 | 135.1761
2015-04-18 | 135.1892
2015-04-24 | 135.1752
2015-04-30 | 135.3038
2015-05-06 | 135.3566
2015-05-12 | 135.1603
2015-05-18 | 135.078
2015-05-24 | 135.18
2015-05-30 | 135.0365
2015-06-05 | 135.1706
2015-06-11 | 135.2884
2015-06-17 | 135.3284
2015-06-23 | 135.2644
2015-06-29 | 135.3803
2015-07-05 | 135.3754
2015-07-11 | 135.4352
2015-07-17 | 135.6302
2015-07-23 | 135.1825
2015-07-29 | 135.3576
2015-08-04 | 135.3909
2015-08-10 | 135.4013
2015-08-16 | 135.4958
2015-08-22 | 135.4197
2015-08-28 | 135.2755
2015-09-03 | 135.506
2015-09-09 | 135.5438
2015-09-15 | 135.4825
2015-09-21 | 135.6096
2015-09-27 | 135.4635
2015-10-03 | 135.2128
2015-10-09 | 135.3872
2015-10-15 | 135.5019
2015-10-21 | 135.5808
2015-10-27 | 135.5144
2015-11-02 | 135.5974
2015-11-08 | 135.6012
2015-11-14 | 135.5755
2015-11-20 | 135.5927
2015-11-26 | 135.4522
2015-12-02 | 135.4898
2015-12-08 | 135.5492
2015-12-14 | 135.4051
2015-12-20 | 135.6911
2015-12-26 | 135.5035
2016-01-01 | 135.6385
2016-01-07 | 135.6247
2016-01-13 | 135.4605
2016-01-19 | 135.6682
2016-01-25 | 135.5189
2016-01-31 | 135.255
2016-02-06 | 135.6458
2016-02-12 | 135.4892
2016-02-18 | 135.5917
2016-02-24 | 135.8704
2016-03-01 | 135.7807
2016-03-07 | 135.9702
2016-03-13 | 135.8264
2016-03-19 | 135.675
2016-03-25 | 135.7543
2016-03-31 | 135.816
2016-04-06 | 135.7134
2016-04-12 | 135.7238
2016-04-18 | 135.7171
2016-04-24 | 135.7103
2016-04-30 | 135.7886
2016-05-06 | 135.8943
2016-05-12 | 135.9172
2016-05-18 | 135.9583
2016-05-24 | 135.9034
2016-05-30 | 135.9416
2016-06-05 | 135.8365
2016-06-11 | 135.8279
2016-06-17 | 135.8625

## 2c Solution
Description: Blank filters are weighed on a scale. The scientist wants to use an average of the last 80 filters as the blank reading for all filters of the same type (same manufacturer and part number). The scientist knows there was a new lot opened when they initiated research on March 1, 2015 and there are 10 filters per lot. The scale's calibration certificate is provided below. What is the expanded measurement uncertainty with a confidence level of 95% for using an average blank value?
Example Application: PM10 particulate collection filters.
Food for Thought: (1) What is wrong with this uncertainty analysis? Hint: Check the distribution. (2) Can the scientist use the average approach in this case?
Answer: (1) The distribution is rectangual. Upon further investigation you notice the variation within lots are normally distributed and the averages of the lots are normally distributed. Non-normality can be an indicator of systematic errors in analysis. (2) Not, but an option could be to take the first filter from each lot and weigh it for the blank values for the lot and take the maximum standard deviation found for the 8 different lots to represent the uncertainty.
Data:
Date | Filter Pre-weight (mg) |  | From Lot Data Set: |  | From Whole Data Set: |  | From the Scale Calibration Certificate:
2015-03-01 | 135.1632 | 1 | average | 135.1251 | average | 135.4925
2015-03-07 | 135.0692 | 2 | standard deviation | 0.096018 | standard deviation | 0.25798 | Expanded Uncertainty | 0.0058
2015-03-13 | 135.1123 | 3 | count | 10 | count | 80 | Coverage Factor | 2
2015-03-19 | 135.2624 | 4 |  |  |  |  | Calibration Uncertainty | 0.0029
2015-03-25 | 135.0345 | 5 |  |  | maximum lot standard deviation | 0.12854
2015-03-31 | 135.1475 | 6
2015-04-06 | 134.9215 | 7 |  |  |  |  | Uncertainty Components | n | Uncertainty (mg)
2015-04-12 | 135.1761 | 8 |  |  | Bin | Frequency | Calibration Uncertainty | 100 | 0.0029
2015-04-18 | 135.1892 | 9 |  |  | 135.1126 | 1 | Repeatability | 10 | 0.12854
2015-04-24 | 135.1752 | 10 |  |  | 135.2067 | 9 | Combined Uncertainty |  | 0.12857
2015-04-30 | 135.3038 | 1 | average | 135.2167 | 135.3009 | 10 | Degrees of Freedom | 9.0092
2015-05-06 | 135.3566 | 2 | standard deviation | 0.10823 | 135.395 | 10 | Expansion Factor | 2.2622
2015-05-12 | 135.1603 | 3 | count | 10 | 135.4891 | 9 | Expanded Uncertainty |  | 0.29085
2015-05-18 | 135.078 | 4 |  |  | 135.5833 | 11
2015-05-24 | 135.18 | 5 |  |  | 135.6774 | 8
2015-05-30 | 135.0365 | 6 |  |  | 135.7716 | 11
2015-06-05 | 135.1706 | 7 |  |  | More | 10
2015-06-11 | 135.2884 | 8
2015-06-17 | 135.3284 | 9
2015-06-23 | 135.2644 | 10
2015-06-29 | 135.3803 | 1 | average | 135.4069
2015-07-05 | 135.3754 | 2 | standard deviation | 0.1123
2015-07-11 | 135.4352 | 3 | count | 10
2015-07-17 | 135.6302 | 4
2015-07-23 | 135.1825 | 5
2015-07-29 | 135.3576 | 6
2015-08-04 | 135.3909 | 7
2015-08-10 | 135.4013 | 8
2015-08-16 | 135.4958 | 9
2015-08-22 | 135.4197 | 10
2015-08-28 | 135.2755 | 1 | average | 135.4564
2015-09-03 | 135.506 | 2 | standard deviation | 0.12854
2015-09-09 | 135.5438 | 3 | count | 10
2015-09-15 | 135.4825 | 4
2015-09-21 | 135.6096 | 5
2015-09-27 | 135.4635 | 6
2015-10-03 | 135.2128 | 7
2015-10-09 | 135.3872 | 8
2015-10-15 | 135.5019 | 9
2015-10-21 | 135.5808 | 10
2015-10-27 | 135.5144 | 1 | average | 135.5469
2015-11-02 | 135.5974 | 2 | standard deviation | 0.083403
2015-11-08 | 135.6012 | 3 | count | 10
2015-11-14 | 135.5755 | 4
2015-11-20 | 135.5927 | 5
2015-11-26 | 135.4522 | 6
2015-12-02 | 135.4898 | 7
2015-12-08 | 135.5492 | 8
2015-12-14 | 135.4051 | 9
2015-12-20 | 135.6911 | 10
2015-12-26 | 135.5035 | 1 | average | 135.5396
2016-01-01 | 135.6385 | 2 | standard deviation | 0.12431
2016-01-07 | 135.6247 | 3 | count | 10
2016-01-13 | 135.4605 | 4
2016-01-19 | 135.6682 | 5
2016-01-25 | 135.5189 | 6
2016-01-31 | 135.255 | 7
2016-02-06 | 135.6458 | 8
2016-02-12 | 135.4892 | 9
2016-02-18 | 135.5917 | 10
2016-02-24 | 135.8704 | 1 | average | 135.7847
2016-03-01 | 135.7807 | 2 | standard deviation | 0.0887
2016-03-07 | 135.9702 | 3 | count | 10
2016-03-13 | 135.8264 | 4
2016-03-19 | 135.675 | 5
2016-03-25 | 135.7543 | 6
2016-03-31 | 135.816 | 7
2016-04-06 | 135.7134 | 8
2016-04-12 | 135.7238 | 9
2016-04-18 | 135.7171 | 10
2016-04-24 | 135.7103 | 1 | average | 135.8641
2016-04-30 | 135.7886 | 2 | standard deviation | 0.075708
2016-05-06 | 135.8943 | 3 | count | 10
2016-05-12 | 135.9172 | 4
2016-05-18 | 135.9583 | 5
2016-05-24 | 135.9034 | 6
2016-05-30 | 135.9416 | 7
2016-06-05 | 135.8365 | 8
2016-06-11 | 135.8279 | 9
2016-06-17 | 135.8625 | 10

## Exercise 2d
Description: Blank filters are weighed on a scale. The scientist wants to use an average of the last 80 filters as the blank reading for all filters of the same type (same manufacturer and part number). The scale's calibration certificate is provided below. What is the expanded measurement uncertainty with a confidence level of 95% for using an average blank value?
Example Application: PM10 particulate collection filters.
Food for Thought: How can you rationalize taking a data point out of a data set?
Data:
Date | Filter Pre-weight (mg)
2015-03-01 | 135.45
2015-03-07 | 136.3049
2015-03-13 | 135.7995
2015-03-19 | 135.4888
2015-03-25 | 135.3573
2015-03-31 | 135.4784
2015-04-06 | 135.5135
2015-04-12 | 135.5215
2015-04-18 | 135.1755
2015-04-24 | 135.0218
2015-04-30 | 135.2287
2015-05-06 | 135.4849
2015-05-12 | 135.792
2015-05-18 | 134.9838
2015-05-24 | 135.5607
2015-05-30 | 135.65
2015-06-05 | 135.8347
2015-06-11 | 135.5308
2015-06-17 | 135.5665
2015-06-23 | 135.3603
2015-06-29 | 136.1732
2015-07-05 | 135.6407
2015-07-11 | 135.6118
2015-07-17 | 135.8227
2015-07-23 | 135.4522
2015-07-29 | 135.3028
2015-08-04 | 135.3277
2015-08-10 | 135.7973
2015-08-16 | 135.6123
2015-08-22 | 135.3364
2015-08-28 | 135.2869
2015-09-03 | 135.8288
2015-09-09 | 135.4157
2015-09-15 | 135.8036
2015-09-21 | 135.7351
2015-09-27 | 135.4009
2015-10-03 | 135.6633
2015-10-09 | 135.1611
2015-10-15 | 135.6099
2015-10-21 | 135.4695
2015-10-27 | 135.1065
2015-11-02 | 135.7834
2015-11-08 | 135.3693
2015-11-14 | 135.7481
2015-11-20 | 135.1147
2015-11-26 | 135.5721
2015-12-02 | 135.5737
2015-12-08 | 135.4232
2015-12-14 | 135.5323
2015-12-20 | 135.848
2015-12-26 | 135.3074
2016-01-01 | 135.5799
2016-01-07 | 135.338
2016-01-13 | 135.9204
2016-01-19 | 135.3368
2016-01-25 | 135.5622
2016-01-31 | 135.6835
2016-02-06 | 135.4802
2016-02-12 | 135.8103
2016-02-18 | 135.7168
2016-02-24 | 135.558
2016-03-01 | 135.3512
2016-03-07 | 135.449
2016-03-13 | 135.483
2016-03-19 | 135.3704
2016-03-25 | 135.9062
2016-03-31 | 135.2178
2016-04-06 | 135.3024
2016-04-12 | 135.381
2016-04-18 | 135.4394
2016-04-24 | 135.7064
2016-04-30 | 135.8525
2016-05-06 | 135.6559
2016-05-12 | 135.1548
2016-05-18 | 135.3512
2016-05-24 | 135.8562
2016-05-30 | 135.4316
2016-06-05 | 135.8512
2016-06-11 | 135.6148
2016-06-17 | 135.4016

## 2d Solution
Description: Blank filters are weighed on a scale. The scientist wants to use an average of the last 80 filters as the blank reading for all filters of the same type (same manufacturer and part number). The scale's calibration certificate is provided below. What is the expanded measurement uncertainty with a confidence level of 95% for using an average blank value?
Example Application: PM10 particulate collection filters.
Food for Thought: How can you rationalize taking a data point out of a data set?
Answer: Outlier removal. Various methods, but here is one. 1.5 * interquartile range (IQR)
Data: |  | IQR Outlier Removal Analysis |  |  |  | Post Outlier Removal |  | From the Scale Calibration Certificate:
Date | Filter Pre-weight (mg) |  |  | Date | Filter Pre-weight (mg)
2015-03-01 | 135.45 | average | 135.5337 | 2015-03-01 | 135.45 | average | 135.5239 | Expanded Uncertainty | 0.0058
2015-03-07 | 136.3049 | standard deviation | 0.24819 | 2015-03-07 |  | standard deviation | 0.23381 | Coverage Factor | 2
2015-03-13 | 135.7995 | count | 80 | 2015-03-13 | 135.7995 | count | 79 | Calibration Uncertainty | 0.0029
2015-03-19 | 135.4888 | median | 135.5175 | 2015-03-19 | 135.4888 | median | 135.5135
2015-03-25 | 135.3573 | Quartile 1 | 135.3501 | 2015-03-25 | 135.3573 | Quartile 1 | 135.3558
2015-03-31 | 135.4784 | Quartile 3 | 135.6849 | 2015-03-31 | 135.4784 | Quartile 3 | 135.6712 | Uncertainty Components | n | Uncertainty (mg)
2015-04-06 | 135.5135 | IQR | 0.33481 | 2015-04-06 | 135.5135 | IQR | 0.31541 | Calibration Uncertainty | 100 | 0.0029
2015-04-12 | 135.5215 | high outlier | 136.1871 | 2015-04-12 | 135.5215 | high outlier | 136.1443 | Repeatability | 79 | 0.23381
2015-04-18 | 135.1755 | low outlier | 134.8479 | 2015-04-18 | 135.1755 | low outlier | 134.8827 | Combined Uncertainty |  | 0.23383
2015-04-24 | 135.0218 |  |  | 2015-04-24 | 135.0218 |  |  | Degrees of Freedom | 78.024
2015-04-30 | 135.2287 |  |  | 2015-04-30 | 135.2287 |  |  | Expansion Factor | 1.9908
2015-05-06 | 135.4849 |  |  | 2015-05-06 | 135.4849 |  |  | Expanded Uncertainty |  | 0.46551
2015-05-12 | 135.792 |  |  | 2015-05-12 | 135.792
2015-05-18 | 134.9838 |  |  | 2015-05-18 | 134.9838
2015-05-24 | 135.5607 |  |  | 2015-05-24 | 135.5607
2015-05-30 | 135.65 |  |  | 2015-05-30 | 135.65
2015-06-05 | 135.8347 |  |  | 2015-06-05 | 135.8347
2015-06-11 | 135.5308 |  |  | 2015-06-11 | 135.5308
2015-06-17 | 135.5665 |  |  | 2015-06-17 | 135.5665
2015-06-23 | 135.3603 |  |  | 2015-06-23 | 135.3603
2015-06-29 | 136.1732 |  |  | 2015-06-29 | 136.1732
2015-07-05 | 135.6407 |  |  | 2015-07-05 | 135.6407
2015-07-11 | 135.6118 |  |  | 2015-07-11 | 135.6118
2015-07-17 | 135.8227 |  |  | 2015-07-17 | 135.8227
2015-07-23 | 135.4522 |  |  | 2015-07-23 | 135.4522
2015-07-29 | 135.3028 |  |  | 2015-07-29 | 135.3028
2015-08-04 | 135.3277 |  |  | 2015-08-04 | 135.3277
2015-08-10 | 135.7973 |  |  | 2015-08-10 | 135.7973
2015-08-16 | 135.6123 |  |  | 2015-08-16 | 135.6123
2015-08-22 | 135.3364 |  |  | 2015-08-22 | 135.3364
2015-08-28 | 135.2869 |  |  | 2015-08-28 | 135.2869
2015-09-03 | 135.8288 |  |  | 2015-09-03 | 135.8288
2015-09-09 | 135.4157 |  |  | 2015-09-09 | 135.4157
2015-09-15 | 135.8036 |  |  | 2015-09-15 | 135.8036
2015-09-21 | 135.7351 |  |  | 2015-09-21 | 135.7351
2015-09-27 | 135.4009 |  |  | 2015-09-27 | 135.4009
2015-10-03 | 135.6633 |  |  | 2015-10-03 | 135.6633
2015-10-09 | 135.1611 |  |  | 2015-10-09 | 135.1611
2015-10-15 | 135.6099 |  |  | 2015-10-15 | 135.6099
2015-10-21 | 135.4695 |  |  | 2015-10-21 | 135.4695
2015-10-27 | 135.1065 |  |  | 2015-10-27 | 135.1065
2015-11-02 | 135.7834 |  |  | 2015-11-02 | 135.7834
2015-11-08 | 135.3693 |  |  | 2015-11-08 | 135.3693
2015-11-14 | 135.7481 |  |  | 2015-11-14 | 135.7481
2015-11-20 | 135.1147 |  |  | 2015-11-20 | 135.1147
2015-11-26 | 135.5721 |  |  | 2015-11-26 | 135.5721
2015-12-02 | 135.5737 |  |  | 2015-12-02 | 135.5737
2015-12-08 | 135.4232 |  |  | 2015-12-08 | 135.4232
2015-12-14 | 135.5323 |  |  | 2015-12-14 | 135.5323
2015-12-20 | 135.848 |  |  | 2015-12-20 | 135.848
2015-12-26 | 135.3074 |  |  | 2015-12-26 | 135.3074
2016-01-01 | 135.5799 |  |  | 2016-01-01 | 135.5799
2016-01-07 | 135.338 |  |  | 2016-01-07 | 135.338
2016-01-13 | 135.9204 |  |  | 2016-01-13 | 135.9204
2016-01-19 | 135.3368 |  |  | 2016-01-19 | 135.3368
2016-01-25 | 135.5622 |  |  | 2016-01-25 | 135.5622
2016-01-31 | 135.6835 |  |  | 2016-01-31 | 135.6835
2016-02-06 | 135.4802 |  |  | 2016-02-06 | 135.4802
2016-02-12 | 135.8103 |  |  | 2016-02-12 | 135.8103
2016-02-18 | 135.7168 |  |  | 2016-02-18 | 135.7168
2016-02-24 | 135.558 |  |  | 2016-02-24 | 135.558
2016-03-01 | 135.3512 |  |  | 2016-03-01 | 135.3512
2016-03-07 | 135.449 |  |  | 2016-03-07 | 135.449
2016-03-13 | 135.483 |  |  | 2016-03-13 | 135.483
2016-03-19 | 135.3704 |  |  | 2016-03-19 | 135.3704
2016-03-25 | 135.9062 |  |  | 2016-03-25 | 135.9062
2016-03-31 | 135.2178 |  |  | 2016-03-31 | 135.2178
2016-04-06 | 135.3024 |  |  | 2016-04-06 | 135.3024
2016-04-12 | 135.381 |  |  | 2016-04-12 | 135.381
2016-04-18 | 135.4394 |  |  | 2016-04-18 | 135.4394
2016-04-24 | 135.7064 |  |  | 2016-04-24 | 135.7064
2016-04-30 | 135.8525 |  |  | 2016-04-30 | 135.8525
2016-05-06 | 135.6559 |  |  | 2016-05-06 | 135.6559
2016-05-12 | 135.1548 |  |  | 2016-05-12 | 135.1548
2016-05-18 | 135.3512 |  |  | 2016-05-18 | 135.3512
2016-05-24 | 135.8562 |  |  | 2016-05-24 | 135.8562
2016-05-30 | 135.4316 |  |  | 2016-05-30 | 135.4316
2016-06-05 | 135.8512 |  |  | 2016-06-05 | 135.8512
2016-06-11 | 135.6148 |  |  | 2016-06-11 | 135.6148
2016-06-17 | 135.4016 |  |  | 2016-06-17 | 135.4016

## Exercise 2e
Description: The scientist was happy with this analysis from scenario 2a and utilizes the average as his blank reading for the collection on July 26, 2016. Below is his final weight. What is the expanded measurement uncertainty with a confidence level of 95% of the dust weight?
Example Application: PM10 particulate collection filters.
Food for Thought: How could you reduce the uncertainty of the measurement?
Measurement Equation:
Y = dust weight
X2 = final filter weight
Data:
Date | Weigh number | Post-filter weight (mg)
2016-07-26 | 1 | 138.9665
| 2 | 138.9645
| 3 | 138.966
| 4 | 138.9654
| 5 | 138.9647
| 6 | 138.9641
| 7 | 138.9652
| 8 | 138.964
| 9 | 138.9656
| 10 | 138.9652
𝑌=𝑋_2−(𝑋_1 ) ̅
(𝑋_1 ) ̅= estimated initial filter weight

## 2e Solution
Description: The scientist was happy with this analysis from scenario 2a and utilizes the average as his blank reading for the collection on July 26, 2016. Below is his final weight. What is the expanded measurement uncertainty with a confidence level of 95% of the dust weight?
Example Application: PM10 particulate collection filters.
Food for Thought: How could you reduce the uncertainty of the measurement?
Answer: Take an actual average measurement of the intial filter weight. See Exercise 2b for the comparison.
Measurement Equation:
Y = dust weight
X2 = final filter weight
Data: |  |  | Post-filter Weight Analysis |  |  | Pre-filter Weight Analysis (from 2a) |  |  | Mesasurement Uncertainty
Date | Weigh number | Post-filter weight (mg) | average | 138.9651 |  | average | 135.5337 |  | Result | 3.4314 | ± | 0.49417
2016-07-26 | 1 | 138.9665 | standard deviation | 0.00080664 |  | standard deviation | 0.24819
| 2 | 138.9645 | count | 10 |  | count | 80
| 3 | 138.966 | standard error | 0.00025508
| 4 | 138.9654
| 5 | 138.9647 | From the Scale Calibration Certificate: |  |  | From the Scale Calibration Certificate:
| 6 | 138.9641
| 7 | 138.9652 | Expanded Uncertainty | 0.0058 |  | Expanded Uncertainty | 0.0058
| 8 | 138.964 | Coverage Factor | 2 |  | Coverage Factor | 2
| 9 | 138.9656 | Calibration Uncertainty | 0.0029 |  | Calibration Uncertainty | 0.0029
| 10 | 138.9652
|  |  | Uncertainty Components | n | Uncertainty (mg) | Uncertainty Components | n | Uncertainty (mg) | Uncertainty Components | n | Uncertainty (mg)
|  |  | Calibration Uncertainty | 100 | 0.0029 | Calibration Uncertainty | 100 | 0.0029 | Pre-filter Weight Analysis | 79.0216 | 0.24821
|  |  | Repeatability | 10 | 0.00025508 | Repeatability | 80 | 0.24819 | Post-filter Weight Analysis | 100.4717 | 0.0029112
|  |  | Combined Uncertainty |  | 0.0029112 | Combined Uncertainty |  | 0.24821 | Combined Uncertainty |  | 0.24822
|  |  | Degrees of Freedom | 100.4717 |  | Degrees of Freedom | 79.0216 |  | Degrees of Freedom | 78.043
|  |  | Expansion Factor | 1.984 |  | Expansion Factor | 1.9905 |  | Expansion Factor | 1.9908
|  |  | Expanded Uncertainty |  | 0.0057757 | Expanded Uncertainty |  | 0.49404 | Expanded Uncertainty |  | 0.49417
𝑌=𝑋_2−(𝑋_1 ) ̅
(𝑋_1 ) ̅= estimated initial filter weight

## Exercise 2f
Description: The scientist decided to do a comparison in uncertainty analysis between using an average estimate for the blank reading and using the average of the actual blank filter for the collection on July 26, 2016. Below includes both pre and post-field weights. What is the expanded measurement uncertainty with a confidence level of 95% of the dust weight?
Example Application: PM10 particulate collection filters.
Food for Thought: Compare results from 2a and 2b. Is the lowest uncertainty always the best practice?
Measurement Equation:
Y = dust weight
X2 = final filter weight
Data:
Date | Weigh number | Pre-filter weight (mg) | Date | Weigh number | Post-filter weight (mg)
2016-07-20 | 1 | 135.433 | 2016-07-26 | 1 | 138.9665
| 2 | 135.4319 |  | 2 | 138.9645
| 3 | 135.4316 |  | 3 | 138.966
| 4 | 135.4322 |  | 4 | 138.9654
| 5 | 135.4314 |  | 5 | 138.9647
| 6 | 135.4305 |  | 6 | 138.9641
| 7 | 135.4306 |  | 7 | 138.9652
| 8 | 135.4323 |  | 8 | 138.964
| 9 | 135.4332 |  | 9 | 138.9656
| 10 | 135.4318 |  | 10 | 138.9652
𝑌=𝑋_2−(𝑋_1 ) ̅
(𝑋_1 ) ̅= estimated initial filter weight

## 2f Solution
Description: The scientist decided to do a comparison in uncertainty analysis between using an average estimate for the blank reading and using the average of the actual blank filter for the collection on July 26, 2016. Below includes both pre and post-field weights. What is the expanded measurement uncertainty with a confidence level of 95% of the dust weight?
Example Application: PM10 particulate collection filters.
Food for Thought: Compare results from 2a and 2b. Is the lowest uncertainty always the best practice?
Answer: Depends on the science question requirement. WMO guidelines sometimes has quoted uncertainties for measurements. If WMO requirement for TSP is +/- 1 mg, using a more efficient method still answers the same question.
Measurement Equation:
Y = dust weight
X2 = final filter weight
Data:
Date | Weigh number | Pre-filter weight (mg) | Date | Weigh number | Post-filter weight (mg)
2016-07-20 | 1 | 135.433 | 2016-07-26 | 1 | 138.9665
| 2 | 135.4319 |  | 2 | 138.9645
| 3 | 135.4316 |  | 3 | 138.966
| 4 | 135.4322 |  | 4 | 138.9654
| 5 | 135.4314 |  | 5 | 138.9647
| 6 | 135.4305 |  | 6 | 138.9641
| 7 | 135.4306 |  | 7 | 138.9652
| 8 | 135.4323 |  | 8 | 138.964
| 9 | 135.4332 |  | 9 | 138.9656
| 10 | 135.4318 |  | 10 | 138.9652
Mesasurement Uncertainty
| average | 135.4318 |  | average | 138.9651 | Actual Blank Measurements Result | 3.5333 | ± | 0.008122
| standard deviation | 0.00088976 |  | standard deviation | 0.00080664 | Estimated Blank Measurements Result | 3.4314 | ± | 0.49417
| count | 10 |  | count | 10
| standard error | 0.00028137 |  | standard error | 0.00025508
From the Scale Calibration Certificate: |  |  | From the Scale Calibration Certificate:
Expanded Uncertainty | 0.0058 |  | Expanded Uncertainty | 0.0058
Coverage Factor | 2 |  | Coverage Factor | 2
Calibration Uncertainty | 0.0029 |  | Calibration Uncertainty | 0.0029
Uncertainty Components | n | Uncertainty (mg) | Uncertainty Components | n | Uncertainty (mg) | Uncertainty Components | n | Uncertainty (mg)
Calibration Uncertainty | 100 | 0.0029 | Calibration Uncertainty | 100 | 0.0029 | Pre-filter Weight Analysis | 100.4717 | 0.0029112
Repeatability | 10 | 0.00028137 | Repeatability | 10 | 0.00025508 | Post-filter Weight Analysis | 100.7744 | 0.0029136
Combined Uncertainty |  | 0.0029136 | Combined Uncertainty |  | 0.0029112 | Combined Uncertainty |  | 0.0041188
Degrees of Freedom | 100.7744 |  | Degrees of Freedom | 100.4717 |  | Degrees of Freedom | 199.246
Expansion Factor | 1.984 |  | Expansion Factor | 1.984 |  | Expansion Factor | 1.972
Expanded Uncertainty |  | 0.0057805 | Expanded Uncertainty |  | 0.0057757 | Expanded Uncertainty |  | 0.008122
𝑌=𝑋_2−(𝑋_1 ) ̅
(𝑋_1 ) ̅= estimated initial filter weight